# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Acquisition of medical equipment	p 1 /…

Publication reference : RORS9/GHPD Smederevo/TD6

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

Impotant Note:

- Tenderer should fill data only in tables for lots he wants to apply and other tables should be deleted.

- For the purpose of evaluation tenderers should provide Manufacturer Catalogue of the each device with marked minimum technical features.

- Tenderer should be aware that delivery of medical equipment should be accompanied with ALIMS certificate (Marketing certificate issued by Serbian Medicines and Medical Devices Agency) in line with Serbian legislative.

LOT 1 EMG

| 1. Lot 1 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L1.1 | EMG 1 set consisting of: Acquisition PC unit - 1 unit - CPU – Minimum Intel I5 or equivalent - RAM memory – Minimum 4 GB RAM DDR4 - Hard Drive – Minimum 500 GB - Professional EMG Sound speakers - Monitor – Minimum 21.5” LCD Full HD Resolution (1920x1080) - Laser printer black and white – Minimal resolution 600x600 dpi - DVD-RW - Operating system – Microsoft Windows 7 Professional or equivalent - Microsoft Office 2016 or equivalent Headbox unit - 1 unit - Trolley for placement of all devices - Isolation transformer – Minimum 500W - Headbox unit - EMG Channels – minimum 2 - Input impedance - >100MΩ - Noise - <1µV r.m.s. - Notch filter – 50/60Hz - Sampling rate – 50KHz per chanel - Sampling resolution – Minimum 16 bit - Sensitivity - 0,1 µV/div to 10 mV/div - Impedance messurment – 0-40KΩ - Power supply – DC & Battery power - Electric stimulation frequency – 0-100Hz - Stimulus Series – Minimum 16 impulses - Pulse width – 0.05ms – 1ms - Gap – 0-12ms - Calibrated headphones - Left, right or Bi-auricolar stimulus presentation - Click, tone or burst stimulation - Visual stimulation by LED Goggles EMG software - 1 set - Patient data base - Patient filtering with advanced search and display features. - Microsoft SQL Data base unique for EEG/EMG/EP. - Automatic display of patient data in report. - Hospital networks by means of support for standard hospital protocols (HL7). - Template based reports. - EMG tests - NEEDLE EMG- spontaneous, voluntary and max. activity - MNCV - SNCV - F-Wave - H-Reflex - Repetitive Nerve Stimulation - SSR - Blink reflex - MUAP - EP tests - SEP - AEP - Option to expand with EP maps - Option to expand with EP grand average - Option to expand with EEG |  |  |  |

LOT 2 EEG

| 1. Lot 2 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L2.1 | EEG - 1 set consisting of: Acquisition PC unit - 1unit - CPU - Minimum Intel I5 or equivalent - RAM memory – Minimum 4 GB RAM DDR4 - Hard Drive – Minimum 500GB - Monitor – Minimum 21.5” LCD Full HD Resolution (1920x1080) - Operating system – Microsoft Windows 7 Professional or better/equivalent - Trolley for placement of acquisition unit - Isolation transformer – Minimum 500W Amplifier unit - 1 unit - Minimal number of Channels – 48 - Sampling rate – minimum 256 Hz - Input impedance - >1GΩ - A/D Conversion – Minimum 24 bit - Impedance check via integrated LED diodes on headbox - Oxymeter channel integrated – Minimum 1 - Patient marker channel – Minimum 1 - Separate stand for amplifier - Notch Filter – 50/60Hz Flash Stimulator - 1 set - Separate stand for flash stimulator - Automatic or manual stimulation protocols EEG Software - 1 set 1.Patient date base - Patient filtering with advanced search and display features. - Microsoft SQL Data base unique for EEG/EMG/EP. - Automatic display of patient data in report. - Hospital networks by means of support for standard hospital protocols (HL7). - Template based reports. - Possibility to create CD/DVD/USB containing EEG readable on any windows compatible PC. - Option to export or import EEG data and EDF. 2. Acquisition - Software for EEG acquisition - Software for EEG reading - Software for EEG reading in real time during acquisition. - Measure cursor with spectral analysis. - Shortcut keys for user defined annotations. - Chronometer for HV and post-HV time. - Software impedance check. - Different colours for each channel. - Possibility of showing annotations and event with automatic synchronization of EEG of a specific event. - Dedicated montage as recorded. - Minimum number of 20 possible montages. - Automatic sliding through EEG trace with selectable speed. - Analysis of fuses, crisis, spikes - Analysis and display of trends - Option to expand on EMG and/or EP - Option to combine with ambulatory EEG - Option to combine with video EEG 3. Software for brain maping - Two dimensional representation on a rotating spherical map - Play function with automatic forward and backward time moving - Possibility to integrate image in report Review PC unit - 1 unit - CPU – Minimum Intel I5 or better/equivalent - RAM memory – Minimum 4 GB RAM DDR4 - Hard Drive – Minimum 500GB - Monitor – Minimum 21.5” LCD Full HD Resolution (1920x1080) - Laser printer black and white – Minimal resolution 600x600dpi - DVD-RW - Operating system – Microsoft Windows 7 Professional or better/equivalent - Microsoft Office 2016 or better/equivalent - Trolley for placement of review unit - Isolation transformer – Minimum 500W Review station - 1 set 1. Review software - Software for EEG reading. - Software for EEG reading in real time during acquisition. - Measure cursor with spectral analysis. - Shortcut keys for user defined annotations. - Chronometer for automatic event insertion. - Software impedance check. - Different colours for each channel. - Possibility of showing annotations and event with automatic synchronization of EEG of a specific event. - Dedicated montage as recorded. - Minimum number of possible 20 montages. - Automatic sliding through EEG trace with selectable speed. - Analysis of fuses, crisis, spikes - Analysis and display of trends 2. Software for brain mapping - Two dimensional representation on a rotating spherical map - Play function with automatic forward and backward time moving - Possibility to integrate image in report 3. Patient data base - Patient filtering with advanced search and display features. - Microsoft SQL Data base unique for EEG/EMG/EP. - Automatic display of patient data in report. - Hospital networks by means of support for standard hospital protocols (HL7). - Template based reports. - Possibility to create CD/DVD/USB containing EEG readable on any windows compatible PC. - Option to export or import EEG data in EDF. |  |  |  |

LOT 3 ULTRASOUND system

| 1. Lot 3 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L3.1 | ULTRASOUND system - 1 unit General features - Stationary, not portable, multi-purpose Ultrasound Diagnostic Systems which meets various needs in ultrasound examination - Multi processing high-speed digital beam former with min. 12-bit A/D converter - Min. reception frequency range 1.0-18.0 MHz - System Dynamic Range min. 270dB - Minimum 300.000 System Processing Channels - Min. 3 active probe ports for electronic scanning - LED monitor min. 17 inch, resolution min. 1280 x 1024 - Operation panel with min. 10'' TFT Touch Screen panel - Operation panel height adjustment and swivel - Alphanumeric keyboard unit which can be moved under operational panel for efficient space use - It is possible to assign measuring functions to the alphabet keys on the keyboard - Minimum depth range selection in B mode from 0.8 to 40cm. - B/W thermal printer Operating modes - B, M-mode, Color Doppler, Power Doppler, Pulsed Wave Doppler with possibility to displays Doppler Spectrums of two different sample points simultaneously, Triplex mode - High resolution color Doppler technique that can display blood flow information in a high spatial and temporal resolution for visualization of fine vessels - Continuous Wave Doppler available on linear probe - Trapezoidal scan for linear probe - Possibility to upgrade system to following advanced features: assessment of liver fibrosis based on elastography, program for early assessment of atherosclerosis and speckle tracking for cardiology examinations Other features Auto angle correction in PW Doppler - Spatial compound imaging - Real-time Doppler Auto Trace - Needle emphasis function for better visualization of biopsy needle - Auto-optimizer: Optimization of gain and sound velocity in B mode and Optimization of gain, velocity range and baseline shift in PW mode - Imaging technique that generate optimum image and reduces speckle noise by changing filter characteristic depending of input signal - It is possible to printout report including ultrasound images directly with an external PC printer Data management - DICOM network communication software: Storage, Transfer and Print - HDD Storage capacity at min. 400GB - Archiving in DICOM, JPEG or TIFF and AVI format - Min. 5 USB connectors Probes - Offered probes are manufactured in Single crystal, Multi-Layered Crystal or Matrix technology - System is compatible with Phased array, TEE, Laparoscopic and Micro convex probes - Linear probe with frequency range from 2-12 (±1) MHz, width between 35 and 40mm, for vascular and soft tissue examinations - Convex abdominal probe with frequency range from 1-5 (±1) MHz |  |  |  |

LOT 4 MPD Equipment

| 1. Lot 4 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L4.1 | 3-Channel ECG Device with Display – 3pcs. Description of the required feature - Portable 3-ch ECG device - Simultaneous of 12-lead ECG acquisition at the backlit display - Diagonal size of the LCD screen: minumum 4.5" - Integrated thermal recorder with 3 simultaneous ECG traces on the paper width min. 60mm - Button for quick pull-out of the paper from the recorder - Manual or automatic ECG recording mode - Periodic recording of the entire ECG writing with adjustable repetition interval - Recording of ECG rhythm in the duration of min. 60s - Automatic analysis of ECG recordings - Acquisition rate of ECG signal: minimum 8000 samples/second - Analysis rate of ECG signal: minimum 500 samples/second - CMRR faktor: minimum 105dB - Input impedance: minimum 50MOhm - Electrode offset tolerance: minimum ±550mV - EMG filter 25/35Hz, AC filter 50/60Hz and high frequency filter 75/100/150Hz - Integrated foiled keyboard - Key for the direct printing of a copy of the last completely recorded ECG analysis - AC and battery power supply (minimum 120min) with integrated charger - Integrated battery charger - Internal memory for minimum 40 ECG recordings - Defibrillation proof - Weight without battery: maximum 2kg Required accessories for each unit: - ECG cable 10 leads, 1 pcs - Precordial electrodes and clips (6+4), 1 set - Accumulator battery, 1 pcs - Thermal paper, 10 pcs - User manual, 1 pcs |  |  |  |
| L4.2 | Defibrillator biphasic, portable – 1 pcs - Portable biphasic defibrillator for external defibrillation, synchronous cardioversion and AED (automatic external defibrillation) with integrated carrying handle - Color TFT LCD display, diagonal measurement min. 6.5" capable for min. 4 waveforms and 5 numeric parameters - External defibrillation paddles with removable adult size - exposing pediatric size - Charge and shock button needs to be integrated in external paddles, as well as the LED indicator in 3 or more levels for contact quality with the patient’s skin - Battery capacity sufficient for min. 100 defibrillations with 270J or minimum 3 hours of continuous ECG monitoring - Battery charging time: max. 3h - Defibrillation energy selectable: from 2J to 270J or higher, biphasic - Energy charging time from 0 to 200J: max 4 s, powered by AC or powered by battery - Fast ECG base line recovery after defibrillation: max. 3 s - Automatic detection of R peak during synchronized cardio version - Automatic internal energy discharge after predefined time - Automatic daily battery test and defibrillator self check, if powered by AC, with results storing - Visual indicator for self test result (operational or not) at front panel - Visual and sound indication of energy charged - Adjustable alarm limits for HR, SpO2, EtCO2 and apnea time - Automatic detection of pacemaker signal - Automatic arrhythmia detection and storing, with sound and visual alarms - Automatic defibrillation reports, with storing and recording from memory. Defibrillation report ECG waveform length min 20s (10s before and 10s after defibrillation) - Automatic arrhythmia storing, with numeric parameters and printing reports - Onscreen guide with instructions for resolving technical alarms - Integrated 2-ch thermal recorder - Operational temperature: -5 to +45°C or better - Dust and moisture protection: min. IP44 - Vibration protection: EN1789 or better - Defibrillator needs to fulfill helicopter transport standard - Weight including battery and paddles: max 7.5 kg - Signal monitoring function: ECG and SpO2. - Upgradeable for monitoring of NIBP, SpO2 and EtCO2 (intubated and nonintubated). - Upgradeable with bed mount hooks Required accessories: - ECG cable with 3 leads, 1 pcs - Cable with the sensor for pulseoximetry for adults, 1set - Acumulator battery, pcs - Thermal paper, 10 pcs - User manual, 1 pcs |  |  |  |
| L4.3 | Laryngoscope set – 2 pcs - Fiberoptic laringoscope with LED diode 3.5V, located in the laryngoscope handle - Blades autoclavable at least 4000 times - Autoclavable handle with the battery 3.5V at 134°C - Illumination strength 4.000 lux or more - Handle manufactured in accordance with ISO 7376 standard or equivalent - Paediatric straight blade Size 1 - Mackintosh blade Size 2 - Mackintosh blade Size 3 - Mackintosh blade Size 4 - Case for the handle and the blades - Table charger with provision for simultaneous charging of two handles |  |  |  |
| L4.4 | Blood pressure gauge – 10 pcs - BHS validated lifelong measuring technology - Special hardened copper-beryllium membrane resistent to aging - Able to withstand pressure up to 600 mm Hg - Aluminum scale from 0 to 300 mm Hg - Maximum error tolerance +/- 2 mm Hg - Valve basis and air channel with precise metal valve - Made of metal with aluminum hausing - Big manometer for easy reading. Scale min Ø 64 mm - To be submitted: Certificate of the measuring device type issued by the Directorate of Mesures and Precious Metals |  |  |  |
| L4.5 | Stetoscope – 10 pcs - Double head made of stainless steel or aluminum - Membrane for precise hearing Ø 44 mm with particularly flat and soft surface which is skin adaptible and not cold to touch - Anatomically shaped phones with integrated multipositionable spring - Extra soft, changeable and revolvable ear olives - Y hose made of PVC for the protection from external sounds - Length 74 cm - To be submitted with a pair of spare ear olives, additional membrane and name plate |  |  |  |
| L4.6 | Neurological hammer – 3 pcs - Neurological hammer type BUCK - Length 18cm - Needle at the top of the handle - Brush at the foot of the handle |  |  |  |