# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Medical equipment	p 1 /…

Publication reference: 5/RORS-9/2017

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

NOTE:

Unless otherwise specified, the requirements in these Technical Specifications are presented as a minimum standard which the offered goods must meet. The offers that will not satisfy completely and correctly the minimum requirements mentioned below will be disqualified.

When brand names/specifications referring to or describing products of a given brand or origin are used in the technical specifications, they are used in descriptive purposes only or should be interpreted as equivalent.

| 1. | Ultrasound system – 1 piece Doppler colour ultrasound system with 4 neurone transducer. Minimal requirements: The command console should be movable on wheels and horizontal. The console should be movable independent of the monitor. Nominal power consumption maximum: 1.3 kVA LCD monitor minim 21`` with possibility of horizontal rotation min ± 80°, angle of visibility ± 175⁰ Integrated DVD-RW module Integrated min. 12`` touch-screen monitor in the console The dynamic range of the system must be at least 200 dB The framerate should be at least 900 frames/sec High definition zoom Internal HDD with a capacity at laeast 1.5 Tb Termic printer The patient information database should be exportable in the following formats: DICOM, JPG and AVI Minim 4 USB ports Generic measurements: distances, areas, volumes, speeds, time Doppler calculations for speed, PS, ED, RI, PI, HR, TAMAX, TAMEAN |  |  |  |
| --- | --- | --- | --- | --- |
|  | The system must be able to work with 3 active modes at the same time (triplex): - Doppler pulse mode – PRF between 150Hz and 48000 Hz - Doppler colour mode- 0,01 cm/s and 4m/s - Power Doppler mode - PRF between 100Hz and 19000 Hz - Higher harmonic mode - Image processing mode in compounding technique with at least 11 composing angles - Highlighting vascularization and microvascularization mode. Included probes: Linear transducer with a min. frequency between 4MHz and 12 MHz, matrix technology with at least 190 crystals, working in compounding technique with image processing from at least 11 angles, minimum 2 selectable harmonic frequencies, at least 3 frequencies in 2D mode, at least 4 Doppler frequencies, penetration depth: minimum 12 mm, Multifrequency sectorial transducer (phased array), selectable in the minimum frequency band from 1.5MHz to 4MHz, selection of minimum 3 frequencies in 2D, selecting at least 4 harmonic frequencies, at least 110 items, minimum 85⁰ viewing field, fingerprint size: max. 2 cm, penetration depth: minimum 30cm Continuous Doppler transducer, 5 MHz frequency, at least 2 items Continuous Doppler transducer, 2 MHz frequency, at least 2 items 24 months warranty |  |  |  |
| 2. | Diagnostic portable ultrasound systems – 2 pieces Minimal requirements: Ready to use within 35 seconds of pressing the on / off button Monitor minimum 15`` with a resolution at least 1280x768 pixels, angle of visibility ± 175⁰ Command panel should be protected with a transparent film Battery autonomy: minimum 50 minutes ± 10% Weight (including batteries), without trolley: max. 7.5Kg. Internal storage capacity of at least 160GB Storage capacity: minimum 25,000 images CINE memory: up to 1200 images Wireless connection Probes: Linear transducer with a min. frequency between 4MHz and 11 MHz, at least 128 items, fingerprint size: 35-50 mm, minimum 3 selectable harmonic frequencies, at least 3 frequencies in 2D mode, at least 3 Doppler frequencies, penetration depth: minimum 10 cm Multifrequency sectorial transducer (phased array), selectable in the minimum frequency band from 1.5MHz to 4MHz, selection of minimum 3 frequencies in 2D, selecting at least 3 harmonic frequencies, at least 64 items, minimum 90⁰ viewing field, fingerprint size: max. 2 cm, penetration depth: minimum 30cm |  |  |  |
|  | Video printer black/white Specially dedicated trolley featuring black and white digital printer with at least 3 active port connectors for 2D image transducers Technical and functional parameters: Work modes - 2D, M, Doppler Color, Doppler Power, Spectral Pulse and Continuous Doppler Minimum 5,000 processing channels The dynamic range of the system must be at least 200 dB Zoom: minimum X 8 Digital signal processing to capture images up to a minimum of 1200 frames per second The equipment must have at least the following type of probes: convex, linear, transtoracic phase, endocavitary microconvex. Image performance specifications: Multi-angle 2D image modeling to improve spatial resolution (Mod Compound), selectable minimum up to 7 compression angles, active with linear and convex type transducers Harmonic mode: frequent harmonics in the phase inversion technique, superior harmonics, active with linear, convex, and "phased array", up to 4 selectable harmonic frequencies Automatic 2D image optimization mode by pressing a single button Spectral Pulse Doppler Module: minimum: ± 340 cm/s at 0 ° correction of angle, PRF between 250Hz and 14000 Hz, selectable minimum 6 wall filters, angle correction, for vascular applications in the minimum range of 0˚ - 85˚, gate size: adjustable minimum from 1 mm to 20 mm |  |  |  |
|  | Continuous spectral Doppler module: active on phased array transducers, minimum from ± 30cm/s to ± 620cm/s, PRF range between: 1.6KHz - 19KHz, selectable minimum 6 wall filters Doppler color module: minim ±240 cm/s, PRF between 250Hz and 14000 Hz, left/right inclinable color Doppler window for linear transducers, up to 3 Doppler selectable transmission frequencies, image acquisition frequency minimum up to 160 images/second, selecting at least 8 speed maps Doppler Power module: minimum 2 frequencies in the transmission, selectable by the user, PRF between 250Hz and 14000 Hz, Doppler power acquisition frequency: up to 160 images/second, left/right inclinable color Doppler window for linear transducers, at least 8 selectable maps in Doppler power. |  |  |  |