# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Acquisition of medical equipment	p 1 /…

Publication reference : RORS9/GHPD Smederevo/TD6-3

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

Impotant Note:

- Tenderer should fill data only in tables for lots he wants to apply and other tables should be deleted.

- For the purpose of evaluation tenderers should provide Manufacturer Catalogue of the each device with marked minimum technical features.

- Tenderer should be aware that delivery of medical equipment should be accompanied with ALIMS certificate (Marketing certificate issued by Serbian Medicines and Medical Devices Agency) in line with Serbian legislative.

LOT 1 EMG

| 1. Lot 1 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L1.1 | 3-channel EMNG system 1 set consisting of: Computer - Brand name - Processor: Intel i5 min. or equivalent - RAM: 4 GB min. - HDD: 500 GB min. - Laser printer: 600 dpi min. - System cart - Microsoft Windows 7 Professional or equivalent - Microsoft Office 2016 or equivalent - Monitor: 21.5" wide screen Standard software - mandatory - Monitoring - Back/forth moving - Offline averaging - Signal amplifier - EMNG signal quality measuring - Signal smoothing - Aftermeasure filtering - Adjustable automatic stimulation - Real time report creating - Table and numeric report - Auto or manual waveform transfer to report - Wrong result warning if exist in list Test type - mandatory - Motor nerve conduction (MNC) - Sensor nerve conduction (SNC) - Combined sensory index - Motor and sensor nerve combined conductivity - Inching studies - Live play - F wave - H reflex - Blink reflex - Needle EMG - Sympatic skin response (SSR) - Parasimpatic skin response - Decrement - RNS (Repetitive nerve stimulation) - R-R interval - Macro EMG - Carpal tunnel - Tremor Quantitative EMG (qEMG) software - mandatory - Interference pattern Analysis (IP) - Turns & Amplitude (TA) - Fast Fourier Transform (FFT) - Expert quantification of the IP Single Fiber EMG software - mandatory - Triggered Single Fiber EMG - Stimulated Single Fiber EMG - Fiber Density Motor unit number index (MUNIX) software - mandatory - Compound muscle action potential (CMAP) Multi-MUP EMG analysis software - mandatory Evoked potentials software - mandatory - SEP (somatosensory evoked potentials) - VEP (visual evoked potentials) - AEP (auditivni evocirani potencijali) GDT interface software - mandatory Interference Pattern Analysis (TA) software - mandatory Reference value software - mandatory Conectivity - HL7 conectivity option - Microsoft SQL database technology Amplifier - RJ45 connection with base unit - Electricaly isolated, min. 3-ch amplifier - 5-pin DIN and 1.5 mm TP (touch proof) connectors for each channel - Input impedance: >1GΩ/25 pF - IMR: >160 dB - CMRR: more than 124 dB - Noise: max. 0.4 μV rms - Lower frequency (-3 dB) min. 0.2 - 5000 Hz - Upper frequency (-3 dB) min. 0.2 - 5000 Hz - Electrode impedance: 1kΩ - 1MΩ - Sensitivity: min. 10nV/D - 100mV/D Aquisition - A-D conversion: min. 24-bit ADC resolution - A-D conversion: min. 48kHz ADC frequency - Duration: min. 0.5ms/div - 5s/div Electric stimulators - Monophasic and biphasic constant curent stimulator - Handheld stimulator with buttons for intensy, stimulation, anode/cathode changing, next test, next page - Repetition: internal. min. 0.06 - 200 pps (Hz) - Output: constant current 0 - 100 mA - Stimulus duration: min. 0.02 ms - 1 ms |  |  |  |

LOT 2 EEG

| 1. Lot 2 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L2.1 | Video EEG system with reader station - 1 set consisting of: Operational System and System Software for Acquisition Station - Microsoft Windows 7 Professional or equivalent - "Microsoft Office 2016 Professional (Word 2016, Excel 2016, Outlook 2016, PowerPoint 2016, Access 2016) or equivalent - Data import/export in EDF and ASCII Acquisition Computer System PC computer with described characteristics: - Hard disc drive: 500 GB min. - Processor: Intel i5 min or equivalent. - DVD-RW Dual Layer Drive - RAM memory: 4GB DDR3 minimum - 21,5" Wide Screen monitor minimum - Mobile chart for acquisition system Photo Stimulator - Noiseless LED photic - Unlimited number of programs for automatic photic operation - Photic is directly connected to the amplifier EEG Amplifier - Total number of input channels: 32 min. - Separate patient event input - Separate SpO2 sensor input - Separate amplifier and photic stand - RJ45 connection with PC EEG software - mandatory - EEG signal reading software - Real time EEG reading during acquisition - Patient and EEG data base managing (ODBC compatible) - Unlimited montage number - Microsoft Word compatible report generating software Brain Mapping software (for Acquisition Station) - mandatory - Amplitude; Band-power (spectral); Coherent, EEG maps: Spectrum, Copper, Gray-scale, Red Blue, Heat Spike & Seizure detection software (for Acquisition Station) - mandatory - Spike and Seizure analysis with Spike and Seizure automatic detection Trend analysis software (for Acquisition Station) - mandatory - Band Power (Total, Absolute, Relative), Generic, Spectral Edge, Entropy, F Ratio, Peak, Envelope, AIE, Spectrogram, Burst Interval Rate, IBI, BR Video EEG software (for Acquisition Station) - mandatory - Synchronized EEG and video/audio recording software Video hardware - Day&night video camera - Audio recording microphone Operational System and System Software for Reader Station - Microsoft Windows 7 Professional or equivalent - Microsoft Office 2016 Professional (Word 2016, Excel 2016, Outlook 2016, PowerPoint 2016, Access 2016) or equivalent - PC computer with described characteristics: - Hard disc drive: 500GB min. - Processor: Intel i5 min. or equivalent - DVD-RW Dual Layer Drive - RAM memory: 4GB DDR3 min. - 21,5" Wide Screen monitor min. - Speakers - Laser printer: black&white, 600dpi min. Video EEG software (for Reader Station) - mandatory Brain Mapping software (for Reader Station) - mandatory - Amplitude; Band-Power (spectral); Coherent. EEG maps: Spectrum, Copper, Gray-scale, Red Blue, Heat Spike & Seizure detection software (for Reader Station) - mandatory - Spike and Seizure analysis with Spike and Seizure automatic detection Trend analysis in all channels software (for Reading Station) - mandatory - Band Power (Total, Absolute, Relative), Generic, Spectral Edge, Entropy, F Ratio, Peak, Envelope, AIE, Spectrogram, Burst Interval rate, IBI, BR |  |  |  |

LOT 3 MPD Equipment

| 1. Lot 3 Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| L3.1 | 3-Channel ECG Device with Display – 3pcs. Description of the required feature - Portable 3-ch ECG device - Simultaneous of 12-lead ECG acquisition at the backlit display - Diagonal size of the LCD screen: minumum 4.5" - Integrated thermal recorder with 3 simultaneous ECG traces on the paper width min. 60mm - Button for quick pull-out of the paper from the recorder - Manual or automatic ECG recording mode - Periodic recording of the entire ECG writing with adjustable repetition interval - Recording of ECG rhythm in the duration of min. 60s - Automatic analysis of ECG recordings - Acquisition rate of ECG signal: minimum 8000 samples/second - Analysis rate of ECG signal: minimum 500 samples/second - CMRR faktor: minimum 105dB - Input impedance: minimum 50MOhm - Electrode offset tolerance: minimum ±550mV - EMG filter 25/35Hz, AC filter 50/60Hz and high frequency filter 75/100/150Hz - Integrated foiled keyboard - Key for the direct printing of a copy of the last completely recorded ECG analysis - AC and battery power supply (minimum 120min) with integrated charger - Integrated battery charger - Internal memory for minimum 40 ECG recordings - Defibrillation proof - Weight without battery: maximum 2kg Required accessories for each unit: - ECG cable 10 leads, 1 pcs - Precordial electrodes and clips (6+4), 1 set - Accumulator battery, 1 pcs - Thermal paper, 10 pcs - User manual, 1 pcs |  |  |  |
| L3.2 | Defibrillator biphasic, portable – 1 pcs - Portable biphasic defibrillator for external defibrillation, synchronous cardioversion and AED (automatic external defibrillation) with integrated carrying handle - Color TFT LCD display, diagonal measurement min. 6.5" capable for min. 4 waveforms and 5 numeric parameters - External defibrillation paddles with removable adult size - exposing pediatric size - Charge and shock button needs to be integrated in external paddles, as well as the LED indicator in 3 or more levels for contact quality with the patient’s skin - Battery capacity sufficient for min. 100 defibrillations with 270J or minimum 3 hours of continuous ECG monitoring - Battery charging time: max. 3h - Defibrillation energy selectable: from 2J to 270J or higher, biphasic - Energy charging time from 0 to 200J: max 4 s, powered by AC or powered by battery - Fast ECG base line recovery after defibrillation: max. 3 s - Automatic detection of R peak during synchronized cardio version - Automatic internal energy discharge after predefined time - Automatic daily battery test and defibrillator self check, if powered by AC, with results storing - Visual indicator for self test result (operational or not) at front panel - Visual and sound indication of energy charged - Adjustable alarm limits for HR, SpO2, EtCO2 and apnea time - Automatic detection of pacemaker signal - Automatic arrhythmia detection and storing, with sound and visual alarms - Automatic defibrillation reports, with storing and recording from memory. Defibrillation report ECG waveform length min 20s (10s before and 10s after defibrillation) - Automatic arrhythmia storing, with numeric parameters and printing reports - Onscreen guide with instructions for resolving technical alarms - Integrated 2-ch thermal recorder - Operational temperature: -5 to +45°C or better - Dust and moisture protection: min. IP44 - Vibration protection: EN1789 or better - Defibrillator needs to fulfill helicopter transport standard - Weight including battery and paddles: max 7.5 kg - Signal monitoring function: ECG and SpO2. - Upgradeable for monitoring of NIBP, SpO2 and EtCO2 (intubated and nonintubated). - Upgradeable with bed mount hooks Required accessories: - ECG cable with 3 leads, 1 pcs - Cable with the sensor for pulseoximetry for adults, 1set - Acumulator battery, pcs - Thermal paper, 10 pcs - User manual, 1 pcs |  |  |  |
| L3.3 | Laryngoscope set – 2 pcs - Fiberoptic laringoscope with LED diode 3.5V, located in the laryngoscope handle - Blades autoclavable at least 4000 times - Autoclavable handle with the battery 3.5V at 134°C - Illumination strength 4.000 lux or more - Handle manufactured in accordance with ISO 7376 standard or equivalent - Paediatric straight blade Size 1 - Mackintosh blade Size 2 - Mackintosh blade Size 3 - Mackintosh blade Size 4 - Case for the handle and the blades - Table charger with provision for simultaneous charging of two handles |  |  |  |
| L3.4 | Blood pressure gauge – 10 pcs - BHS validated lifelong measuring technology - Special hardened copper-beryllium membrane resistent to aging - Able to withstand pressure up to 600 mm Hg - Aluminum scale from 0 to 300 mm Hg - Maximum error tolerance +/- 2 mm Hg - Valve basis and air channel with precise metal valve - Made of metal with aluminum hausing - Big manometer for easy reading. Scale min Ø 64 mm - To be submitted: Certificate of the measuring device type issued by the Directorate of Mesures and Precious Metals |  |  |  |
| L3.5 | Stetoscope – 10 pcs - Double head made of stainless steel or aluminum - Membrane for precise hearing Ø 44 mm with particularly flat and soft surface which is skin adaptible and not cold to touch - Anatomically shaped phones with integrated multipositionable spring - Extra soft, changeable and revolvable ear olives - Y hose made of PVC for the protection from external sounds - Length 74 cm - To be submitted with a pair of spare ear olives, additional membrane and name plate |  |  |  |
| L3.6 | Neurological hammer – 3 pcs - Neurological hammer type BUCK - Length 18cm - Needle at the top of the handle - Brush at the foot of the handle |  |  |  |