# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Supply of laboratory furniture, instrument and other laboratory equipment	p 1 /19

Publication reference : RORS337/TD03

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT1 Supply of laboratory furniture

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Safety cabinet for chemicals for acids and bases Complying with EN 61010-1. Constructed in FE P01 grade, 10/10 thickness, electro-galvanized steel sheet. “Monolithic” design, with adjustable feet at the base to allow the structure to be properly levelled. “Sandwich-wall”, reversible doors must be constructed in the same materials and come complete with handles incorporating safety closures and locks. Doors should be open to 110° which allows internal shelves to be removed without having to incline them. On te fron doors should be multilingual safety signs which indicate the presence of dangerous and toxic products External dimensions: 1150 ±10x610 ±10x1700 ±100 mm (LxWxH) Should be adapted for use with external extraction systems Should be adapted for use with filtration systems NON-SPARKING EXTRACTION SYSTEM in metal box ON / OFF fuses power cable, filter box with carbon filter |  |  |  |

LOT 2 Supply of instrument optical emission spectrometer with inductively compound plasma (ICP-OES)

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Instrument optical emission spectrometer with inductively coupled plasma (ICP-OES) Automatic, simultaneous measuring optical emission spectrometer with inductively coupled plasma ICPOES -Plasma observation in two axis for measurement of trace elements in environmental samples, -Thermally stabilized optical system (+15°C ± 0.5°C), -Possibility of measuring Cl at 134.7nm to K at 766.5nm, -Detectors, linear CCD arrays, fitted in continuous order,-Dynamic range up to 9 orders of magnitude, -Monitoring of all elements spectrum wave lengths, -Automatic background correction, -Automatic profile standardization, -Optical system filled with argon gas, -Frequency: 27.12 MHz, free running type, -RF power output: 0.5 to 2.0 kW , -Power efficiency: >85%, -Power stability <0.1%, -Air cooled (No external cooling required), |  |  |  |
|  | Standard sample introduction system, consisting of Demountable Plasma Torch, Cross flow nebulizer, Miniature Scott (double pass) spray chamber for fast washout, Fast pre-flush and fast rinse function Automatic wavelength standardization with one standard. Whole-spectrum standardization based on all pixels -Enables transfer between individual instruments -Automatic drift control during the measurement process -Recording of complete spectrum or region of interest (with data storage) during sample/standard measurement -Data Reprocessing enabled -Storage of the complete spectrum for every measurement -Analysis with modified peak(area) and/or background settings -Complete lines library (NIST) -PC with appropriate software, -Led 17 inch Monitor, -Color Printer |  |  |  |
|  | Instrument installation and commissioning: The Contractor should provide training for operation the instrument for 3 persons from the Contracting Authority. The training should be organized within maximum 5 days after comissioning. The training should last minimum three consecutive days, starting from 7 AM until 3PM. Upon completion of the training, the Supplier is required to issue training Certificates to each trainee. The Tenderer must be able to arrive within 24 hours if there is a failure of the instrument within the warranty period. |  |  |  |

LOT 3 Supply of Laboratory equipment and laboratory consumables

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Glass reactor 1 pieces The reactor must be made of borsilicate glass, resistant to aggressive chemicals and high temperature. It must have automatic control of temperature and stirring speed. Reactor is consisting of the following items: reaction vessel 100 L with inclined agitator unit, electrical heating mantle, main condenser and support structure. For chemical reaction which happen inside of the reactor it is very important that reactor have inclined agitator. Temperature rang: -20 to +150 °C Pressure rang: -1 to +0,5 bar Reaction Unit must consist of: - Spherical glass vessel, nominal volume 100 l, with: 1 x bottom nozzle for for low dead space bottom outlet valve; 1 x top nozzle DN200 (neck) for vapours; 1 x peripheral nozzle for inclined stirrer assembly; 3 x peripheral nozzles for gas pipe inlet (DN 40), quick release coupling for filling procedures (DN 150), temperature measurement (DN 40) - Stirrer: standard stirrer type, inclined assembly, variable speed 0 -600 rpm, with propeller stirrer element |  |  |  |
|  | - Electric heating mantle: capacity: 6kW, 230/400 V, 3 – phase with 4 heating circuits with thermoelement NiCr-Ni, with control unit including temperature controller installed int he framework and Pt 100 resistance thermometer with transmitter - Reflux condenser: type - coil heat exchanger, glass, exchange area: 1 m2 and heat exchanger with top reducer and with hose connectors on utility side - Accessories must include: low bottom outlet valve; pipe inlet for gas; quick release coupling, couplings with bolts and nuts; gasket made of PTFE; bellows at the vent connection and pipe inlet; vapour piping between the reaction vessel and the condenser; support elements to carry the equipment - The framework must be made of stainless steel - Dimensions of the unit: area approx. 1000 x 1200 mm and height approx. 2400mm - Instrumentation and Control Equipment must include: temperature sensor (PT100) with transmitter and control panel (temperature controller) for heating mantle |  |  |  |
|  | - Process pipework must include: Fort he process piping tubes made of Borosilicate glass 3.3 and/or PFA hoses are used; All valves (flanged type version for better cleaning purposes) int he plant are made of Borosilicate glass 3.3/PTFE; All gaskets made of PTFE ; Flanges up and including DN300 made of stainless steel (QVF SUPRA Line); Flanges over DN300 are epoxy coated cast iron; All bellows / couplings required for interconnecting glass piping to non-glass piping All units, apparatus and glass products comply with the following International Standards: - 2014/68/EU Pressure Equipment Directive - 2006/42/CE Machinery Directive (e.g. for Stirrer Unit/Pump) - 2014/34/EU Explosion Prevention Directive/ATEX - 2014/35/EU Low Voltage Directive - AD 2000Design and safety requirements - DIN EN 1595 Pressure equipment made from borosilicate glass 3.3 - TRGS 727Protection against electrostatic charge - DIN EN 12585 Glass plant, pipeline and fittings - DIN ISO 3585 Properties of borosilicate glass 3.3 - DIN EN National/Europian Standards (where required) - ISO International Standards (where required) |  |  |  |
|  | Corresponding to these Directives and Standards all apparatus and glass products are marked with CE-sign acc. to Pressure Equipment Directive. Requirements of the international Quality Management System according to DIN EN ISO 9001. The supplier should provide delivery, installation and manuals. |  |  |  |
| 2 | Balance for measuring larger samples, from 60 to 65 kg 1 piece Readability (d) min 0,5 g Weighing range (max) 65.000 Taring range (subtractive) max 65.000 Reproducibility min 0,5 g Linearity min ± 1,5 g Smallest piece weight min 0,5 g Adjustment point 20/30/50/60 kg Appropriate for verification max. 80% rel. (non- condensing) Stabilization time (typical) min 3 sec. Allowable ambient temperature min +10 °C + 40 °C Warm-up time max 2 hours Weighing plate stainless steel Data interface yes (RS232) Battery operation with yes |  |  |  |

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 2 | Balance for measuring of smaller samples, from 5 to 6 kg 1 piece Readability (d) min 0,01 g Pan Size (200 ±10 mm x 170 ±10 mm) Internal Calibration AutoCal™ - Automatic Communication RS232 (Included); USB; USB Host Display High resolution full colour VGA touch display In-use cover Included Minimum Weight (USP, 0.1%, typical) 20 g Pan Construction Stainless Steel Repeatability min 0.01 g Stabilization Time minimum 1.5 s Working Environment minimum 50°F – 85°F, 80%RH, non-condensing (10°C – 30°C, 80%RH, non-condensing) Stabilization time min 1.5 sec. Linearity ± min 0.02 g The Tenderer must be able to arrive within 24 hours if there is a failure of the instrument within the warranty period. The supplier should provide delivery, installation and manuals The supplier must have a service on the territory of the Republic of Serbia for all installed parts and equipment |  |  |  |
| 3 | Pump for highly aggressive acid and alkalis 1piece - Barrel pump for highly aggressive acids and alkalis, such as sulphuric acid, nitric acid, hydrofluoric acid, sodium hypochlorite etc. - Motor 230 V - PVDF pump tube - Thread connection 1 1/4” with hose connector 3/4” - 2 m special chemical hose 3/4” - PVDF pump nozzle - PP barrel screw joint - Pump height 17 mWS, pump capacity 80 l/min - Immersion depth cm 100-120 - Tube Ø mm 41-43 Pump for weaker acids 1 piece - Barrel pump for weaker acids, such as hydrochloric acid, phosphoric acid, chromic acid, battery acid, citric acid, etc. - Motor 230 V - Pump tube PP, shaft Hastelloy C - Thread connection 1 1/4” with hose connector 3/4” - 2 m PVC fabric hose 3/4” - PP pump nozzle - PP barrel screw joint - Pump height 19 mWS, pump capacity 85 l/min - Immersion depth cm 100-120 - Tube Ø mm 41-42 The supplier must have a service on the territory of the Republic of Serbia for all installed parts and equipment |  |  |  |
| 4 | Microwave rotor 1 pieces Segmented digestion rotor, with 15 teflon vessels of 100 ml capacity, with standing operating pressure up to 100 bar and temperature up to min 300oC Vessels must have safety shild and a self-ventilating mechanism that releases only excess pressure, which then closes again. Rotor should have the possibility of using quartz and teflon vessels of 5 to 30 ml capacity, used within the primary teflon vessels of the rotor Equal with Milestone ETHOS UP microwave system |  |  |  |
| 5 | Laboratory digestor with set of underground cabinets for chemicals and random surface of monolithic ceramics with ventilation system 1 piece Didestor should be maid according to standard EN14175 Digestor must have vertically raised front panel made of toughened glass, which should have the ability to lift the up to a minimum of 700mm from the work surface, and front panel should have the possibility of locking (fixing at the desired height) Digestor working surface must be of ceramic materials resistant to aggressive substances and temperature up to a minimum of 1200 0C Internal Dimensions (working space should be): min 1300mm x min 650mm External Digestor Dimensions (Width / Depth / Height): Should be not more than 1550x1000x2400mm Digestor mast have conection for elecricity on 220V and electrical safety in accordance with EN-61010-1 |  |  |  |
| 5 | Lighting in the digestor must be minimum 1000 lux, Digestor must have water connection with drainage, one gas connection and minimum four power outlets Digestor have to be supplied with a hose for connection to the drain (longines of min 7m with three knee) and an air blower with the following characteristics. - plastic impellers -made of high resistance materials to chemicals and corrosive vapors performance in accordance with AMCA 210-85 and ISO 5801. engine speeds minimum 1400 rpm power consumption not exceeding 1.4 kW. The supplier should provide delivery, installation and manuals. The supplier must have a service on the territory of the Republic of Serbia for all installed parts and equipment |  |  |  |
| 6 | Standard solutions Standard for Fe 1000ppm of 100 ml in nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Pb 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Zn 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Ni 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Mg 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece |  |  |  |
| 6 | Standard for Al 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Ca 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Cr 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Mo 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Cd 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece |  |  |  |
| 6 | Standard for As 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Se 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Hg 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Bi 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Te 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece |  |  |  |
|  | Standard for Co 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Mn 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Sb 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for Si 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece Standard for P 1000ppm of 100 ml nitric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece |  |  |  |
|  | Mix tandard for REE 1000ppm of 100 ml nitric acid/hydrochloric acid in glass bottle seald and protected from from breaking. Standard solution must have certificate validated not less then 2022 1 piece The supplier should provide Certificate of quality for standard solutions |  |  |  |
| 7 | Chemicals and reagents HCl pa quality in glass bottle seald and protected from breaking in quantitiy of 2.5L 8 pieces HNO3 pa quality in glass bottle seald and protected from breaking in quantitiy of 2.5L 8 pieces HF pa quality in glass bottle seald and protected from breaking in quantitiy of 2.5L 8 pieces H2O2 pa quality in glass bottle seald and protected from breaking in quantitiy of 2.5L 8 pieces The supplier should provide MSDS list about characteristics of chemicals |  |  |  |