# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Supply of Workstation	p 1 /…

Publication reference : RORS9/SBPB Vrsac/TD14

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number 1 | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Workstation Qty: 1 set Hardware specification: - PC with the following min. specification: Intel CPU or appropriate, min. 3.0 GHz, min. 500 GB HDD, min. 8GB DDR3 RAM with graphics card supporting high-resolution monitors - DICOM medical-grade monitor with the following specification: Color (IPS), LED, min. 24 in, resolution: min. 1900 x 1200, contrast: min. 1000:1, brightness: min. 350 cd/m2 - Automated medical grade robot burner - Burning of min. 30 CDs or min. 15 DVDs per hour - Input storage for min. 100 CDs or DVDs - USB 3.0 or 1GB LAN interface - Min. two CD/DVD burners - Automated CD/DVD burning allowed - Automated arm for transport of studies - Burning of study viewers with MPR support included - Automatic printing to the printable side of the CD/DVD of user-defined information - Ink jet printing technology - Burning of multiple views to a single medium allowed Required software specification - Must have unlimited number of software licenses for Multi-planar Reconstruction with min. following functionalities: - Settings of thickness of slice and spacing - MIP and MinIP during the use of MPR - Creation of curved slices (parallel and concentric) allowed - Filming/printing of MPR images - Automatic link, i.e. automatic connection of MPR groups - Min. one concurrent/simultaneous 3D software license with the following functionalities: - Definition of the volume of interest and removal of unnecessary volume areas to get a clearer picture of pathology, - Volumes of interest archiving allowed, - Volume Rendering, - Definition of 3D tissues, - Screenshots allowed for a study within 3D license/window, so that a radiologist may continue with the study diagnostics at a later point, with an unchanged screen appearance (at the same point where he/she interrupted the work), - Additional tools for adjusting the display of the study, i.e. removal of fragments, after the activation of the automatic bone removal tool - Special tool for a single click patient table removal, - Special tool for a single click bone removal - Min. one concurrent/simultaneous license for volumetric comparison with the following functionalities: - Automatic alignment and synchronous examination allowed of the study slices being compared by all three planes (transversal, frontal, sagittal) as well as between different modalities (CT, MR) - Three-dimensional comparison of studies acquired at different periods of time and on different modalities (CT and MR) allowed - An intelligent algorithm of remote loading of exams through slow Internet connections of up to 10Mbit, by the system by which the images/slices the radiologist focuses on as well as the surrounding slices are loaded first during loading and then all the others, too, with a gradual increase in resolution to the original one, so the radiologist could start diagnostics as soon as possible |  |  |  |