# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Supply of Acquisition of laboratory equipment	p 1 / 13

Publication reference: RORS-462/04

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| Lot.1. Laboratory centrifuge |  |  |  |  |
| 1 | Laboratory centrifuge (minimum characteristics requested): Quantity: 1 pcs. |  |  |  |
|  | Multifunctional, with LCD display, rotation speed setting, ability to run with various tube dimensions – from 1.5 to 50 ml – with shutdown protection at unbalanced rotor charge. |  |  |  |
|  | Speed Range: 200 rpm – 6000 rpm; 10 rpm/set |  |  |  |
|  | RCF: up to 4000 × g; 10 × g/set |  |  |  |
|  | Rotor capacity: up to 8 × 15 mL |  |  |  |
|  | Running Time: Quick Spin, 10 sec to 9 hr, or continuous |  |  |  |
|  | Reach maximum acceleration/deceleration: 10 sec. |  |  |  |
|  | Included rotor: 8x15 ml |  |  |  |
|  | Unit cap with mechanical blocking system |  |  |  |
| Lot.2: Thermoreactor with minimum 4 seats for probe digestion |  |  |  |  |
| 1 | Thermoreactor with minimum 4 seats for probe digestion (minimum characteristics requested): Quantity: 1 pcs. |  |  |  |
|  | Electronic temperature control |  |  |  |
|  | Temperature regulation from ambient to 200 °C |  |  |  |
|  | Analysis time can be set from 1 to 160 minutes or continuous |  |  |  |
|  | Processes minimum 4 x 200 ml test tubes for wet digestion |  |  |  |
|  | Temperature stability: ± 0.5 °C |  |  |  |
|  | Temperature homogeneity: ± 0.5 °C |  |  |  |
|  | Temperature precision: ± 1 °C |  |  |  |
|  | Protection to overheating |  |  |  |
|  | Electric heated plate with minimum 4 holes, diameter of 42 mm |  |  |  |
|  | Display for temperature and remaining time visualization |  |  |  |
|  | Visual signals for target temperature reach |  |  |  |
|  | Visual and acoustic signals for cycle over, probe malfunction, broken probe, temperature setting overload. |  |  |  |
|  | minimum 4 air refrigerants |  |  |  |
|  | minimum 4 anti-splash bells |  |  |  |
|  | minimum 4 PTFE sleeves for 29/32 cones |  |  |  |
|  | A set of minimum 3 samples COD diam. 42 x 200mm, 200ml with SN 29/32 cone |  |  |  |
|  | Stainless steel bracket for simultaneous 6 probe manipulation |  |  |  |
|  | Metallic construction, painted for protection against chemical reagents |  |  |  |
| Lot.3: Graphite Atomic Absorption Spectrometer with Zeeman and Deuterium background correction |  |  |  |  |
| 1 | Graphite Atomic Absorption Spectrometer with Zeeman and Deuterium background correction (minimum characteristics requested): Quantity: 1 pcs. |  |  |  |
|  | The spectrometer must contain both atomization techniques (flame and graphite furnace) in a single compact housing, and to be installed with an extractive hood for flame capture. The instrument must have included hydride system, upgradable for coupling to graphite furnace |  |  |  |
|  | Atomization system for the flame and graphite furnace, which will be fully automated and will not require user intervention when changing the atomization technique |  |  |  |
|  | Pneumatic control systems for combustion gases in case of flame analysis and inert gas for graphite furnace use |  |  |  |
|  | Burner for all types of flame with optimum automatic adjustment and automatic flame ignition. The spectrometer must be able to provide free choice of single and double beam mode in flame and hydride mode that allows fully automatic alignment with Deuterium and Zeeman background correction with magnetic field control |  |  |  |
|  | The spectrometer must be able to ensure the automatic change of technique (flame, graphite furnace) directly from software without movements, without mechanical adjustments or manual intervention |  |  |  |
|  | The spectrophotometer must be able to be upgraded further with a direct solids and viscous analysis (including sludge), completely automatic, with a minimum of 40 positions and a micro-balance incorporated into the system |  |  |  |
|  | Wavelength range: minimum 185 - 900 nm |  |  |  |
|  | Detector: standard wide range UV sensitive photomultiplier, automatic wavelength setting with the possibility to select from several characteristic lengths for each element |  |  |  |
|  | A single turret with support for 8 lamps fully automated |  |  |  |
|  | Radiation source: coded cathode lamps for automatic recognition of lamp type, easy for replacement, optimized for user |  |  |  |
|  | OPTICAL SYSTEM: Optical system with Czerny -Tuner monochromator with holographic grating with minimum 1800 lines/ mm and bandwidth between 0.2 and 1.2 nm. Free choice of single beam and double beam mode. Compact optical components for resistance and stability. Optical bench with optical parts to be mounted in a cast for strength and stability with cover to prevent vapours, dust and humidity ingress. High performance optical components to reduce energy loss |  |  |  |
|  | ATOMIZER: The flame atomizer must be nebulizing the mixing / spraying chamber to be of non-corrosive material with overpressure protection. Analyses must be performed using either acetylene-air flame (for temperatures below 2300 ° C) or nitrous / acetylene protoxide flame. The atomizer must be controllable by and through the computer |  |  |  |
|  | GAS CONTROL: Automatic adjustment of the gas mixture (pressure / flow) through the software. Automatic flame recognition and ignition. Automatic flame surveillance. Pressure sensors for all types of gases, capable of stopping the flame from accidentally interrupting the gas supply or the electrical voltage |  |  |  |
|  | BURNER: Automatic recognition of the burner type from the software. Automatic control of safety locking systems for burner, for pressure and flame detection sensors. Automatic shutdown in case of voltage drops, block the flame when the burner is incorrectly mounted. Automatic flame optimization. Oxygen and fuel pressure sensor. Liquid level sensor in the drain vessel. Corrosion resistant nebulizer. Coded burner. Automatic system, with software setting, for mechanical cleaning of the burner head surface and with adjustable frequency of cleaning; cleaning can also be done during working with the flame mode. |  |  |  |
|  | Fully automatic device for the flame system with the following characteristics: - will be operated using the software - will provide a dosage injection, controlled by software, of small sample volumes - will ensure the continuous suction of the washing solution in order to keep the burner temperature constant - will reduce the consumption of reagents and the volume of sample needed for the analysis - will help increase reliability and achieve better reproducibility of measurements - will ensure continuous washing of the mixing chamber and nebulizer. |  |  |  |
|  | GRAFITE FURNACE: Is integrated computer -controlled transversely heated Zeeman graphite furnace. Graphite tube is transversely heated to provide a constant temperature profile over the entire length of the tube and to reduce matrix interferences. Zeeman background correction with variable transverse magnetic field. The capacity of the magnetic field to reach a minimum of 1.0 T Fast heating with linear speed Programmable furnace temperature up to 3000 ° C (in steps of 1 ° C) It is not necessary to change or adjust the autosampler and the graphite furnace when replacing the graphite tubes. Easy replacement of graphite tubes without adjusting operations. Cooling unit for graphite furnace. Built-in color video camera for tracking the pipetting and combustion process. Functions that allow the creation of sequential programs. |  |  |  |
|  | AUTOSAMPLER FOR THE GRAPHITE FURNACE: Autosampler with automatic dilution and intelligence with a dilution factor of at least 1: 500 The autosampler must have all the mechanical and electrical components necessary for taking a correct volume of solution from a selected container, injecting it into the graphite tube and washing the injection needle. The autosampler will contain a tray with min. 100 positions depending on the capacity of the containers, in which will be placed containers made of polypropylene of minimum capacity. 1.5 ml and up to 5 ml used for standard solutions, matrix modifiers, standards and samples. The position of each container will be numbered on the tray. The autosampler will not need realignment after changing the graphite tube with a new one. |  |  |  |
|  | Single element coded lamps, with cavity cathode for 10 elements included (arsenium, lead, mercury, cadmium, manganese, zinc, iron, magnesium, uranium, aluminium) |  |  |  |
|  | 5 graphite tubes |  |  |  |
|  | Compatible portable control unit and printer, with software to start-stop the equipment, to obtain the analysis data and the total control of the equipment, including its diagnosis. |  |  |  |
|  | Software for statistical processing of the results of the analyses performed with the atomic absorption spectrometer, facilities for controlling the quality of the results (eg total system security, data and methods developed through multi-level security options that gives the supervisor the opportunity to decide who can have access to the allowed data) |  |  |  |
|  | Automatic control of all components and accessories of the equipment |  |  |  |
|  | Automatic lamp recognition and automatic setting of all working parameters for the lamp (current, wavelength, slot) |  |  |  |
|  | Construction of calibration curves using functions and the addition method; construction of working methods and their storage |  |  |  |
|  | Reanalyse samples that exceed the highest calibration standard. |  |  |  |
|  | Supplier must assure training course for the personnel designated by the beneficiary. The training will include: - learning the basic functions of the equipment and the operating software (software) for keeping the equipment under control (starting / stopping, performing the verification tests); - knowing the significance and understanding of the role of each working parameter on the stages of carrying out an analysis; - the working method for the flame technique and the graphite furnace technique - training the skills to carry out its internal maintenance operations. The course must be taught in Romanian. |  |  |  |
|  | Warranty term: 24 months after installation and schooling for the instrument. |  |  |  |
| Lot.4: N/C analyzer (TOC, NPOC, POC, TC, TIC, TNb) with NDIR detector for liquid samples |  |  |  |  |
| 1 | N/C analyser (minimum characteristics requested): Quantity: 1 pcs. |  |  |  |
|  | Analyser for the determination of the parameters TOC, TNb , NPOC, TC, TIC and POC in aqueous samples. |  |  |  |
|  | The analyser has to be able to be upgraded for TOC analyses from solids |  |  |  |
|  | Analyser has to be modular instrument |  |  |  |
|  | Instrument must have a portable control unit for control with software specially designed for instrument control |  |  |  |
|  | NDIR detector for C |  |  |  |
|  | Chemiluminiscence detector for N |  |  |  |
|  | Precise CO2 detection using the maintenance-free NDIR detector, equipped with a pulsed, high-power IR radiation source, corrosion-free cell, encapsulated optics and highly sensitive micro detector. |  |  |  |
|  | Instrument principle: catalytic high-temperature combustion up to min. 950 °C |  |  |  |
|  | Calibration with long-term stability, multipoint calibration with a single standard using different injection volumes, automatic selection of the optimum calibration curve for sample measurements |  |  |  |
|  | drying of measuring gas without using desiccants and efficient measuring gas cleaning including monitoring |  |  |  |
|  | Autosampler with a minimum of 47 positions, for samples of 12 and 50 ml, with included magnets and the automatic purging. Autosampler with homogenizing function. |  |  |  |
|  | The method of analysis for TOC and TN for liquid samples has to use temperature of minimum 800 C |  |  |  |
|  | The method of analysis for TOC in solid samples has to use temperature of minimum 1300 C |  |  |  |
|  | Time of analysis for liquid samples is not more than 5 minute for each parameter |  |  |  |
|  | The instrument must allow analysis in a concentration range of not more than 30 000 mg/l carbon and 20 000 mg/l TNb with automatic dilution for liquid samples |  |  |  |
|  | Analyser detection limit: - 4 μg/l C - 0.005 mg/l TNb |  |  |  |
|  | Automatic dilution with dilution factor 1:5 – 1:100 |  |  |  |
|  | Auto zero procedure to be carried out automatically before the analysis |  |  |  |
|  | Analysis of total organic carbon and total nitrogen from a single injection |  |  |  |
|  | Injection of a variable volume of sample in the range 100-1000 μg/l |  |  |  |
|  | Use of the same catalyst for both types of determinations, for the entire field of concentration. |  |  |  |
|  | internal self-control system, electronic gas control, electronic gas loss control, electronic detector control, automatic control and display for maintenance |  |  |  |
|  | Supply gas: oxygen or synthetic air |  |  |  |
|  | The analyser must allow the calibration technique to be performed in several points using a single standard and to perform the automatic selection of the optimum calibration curve for sample analysis. |  |  |  |
|  | Calibration to be stable in the long term |  |  |  |
|  | Reproducibility: 2-3% |  |  |  |
|  | Specialized software for controlling the device must allow: -User defined checklist for quick start of operation -Search, sort and filter functions for methods, calibration and result files -Freely selectable sample types (QC standards, daily factors, various blank types) - sequences of mixed methods can be created - While running a sequence new samples can be added, results can be edited or printed -Sample ID’s can be assigned automatically - multi-point calibration from a single standard - Extension of existing calibration curves by simple adding of calibration points - blank determination and automatic result correction possible |  |  |  |
|  | Warranty term for instrument: 24 months after installation and schooling for the instrument |  |  |  |
|  | Warranty term for NDIR detector: minimum 8 years after installation and schooling for the instrument |  |  |  |
|  | Supplier must assure training course for the personnel designated by the beneficiary. The course must be taught in Romanian. |  |  |  |

Unless otherwise specified, the requirements in this Technical Specifications are presented as minimum standard, which the offered goods must meet. The offers that will not satisfy completely and correctly the minimum requirements mentioned above, will be disqualified.

When brands names/specifications referring to or describing products of a given brand or origin are used in this technical specifications, they are used in descriptive purposes only and should be interpreted as equivalent.