# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Procurement of Medical/Laboratory Equipment II for project "Improvement of the Neoplazic pathology management through the development and implementation of cross border medical excellence center"

Publication reference: RORS-498.TD07.P2

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Microscope for pathology with camera and working station Quantity: 1 set Minimal technical specification: Microscope - Binocular photo tube 50%: 50% with an inclination of min. 30 degrees, adjusting the viewing height to two heights. - 5-position coded objective holder that enables function: • light manager • automatic adjustment of the appropriate measurement ratio - Mechanical stage with drive that extendable by min. 15mm. - two integrated buttons on the stand (captureing images or videos directly on the USB) - Achromatic aplanatic condenser min. 0.9 for objectives 1-100x - white balance filter - Lighting LED min 10W coded, full Koeller, intensity settings on the microscope stand and ECO mode (after 15 minutes of inactivity on the microscope it switches to standby mode). The user can switch to a 35W halogen light on their own - Active light manager that allows, when the light intensity is adjusted for the active objective, the same relative ratio is transmitted for other objectives - Min. Objectives: 5x / 0.15, 10x / 0.25, 20x / 0.45, 40x / 0.65, 63x / 0.80 (or 100x /1.25Oil) - Eyepieces 10x with a field of view of min. 23 mm - Digital microscope colour camera with Ultra HD / min. 4K image resolution, live image min min. 30fps via HDMI, USB 3.0). Sensor size min. 8mm, pixel size min 1.8^m, exposure time min. 1s. The same manufacturer as the microscope. Appropriate adapter and software. - Min.4K monitor min. 27" for direct camera connection Working station: - External Size (W*D*H): Min. 1800*700*1900mm - Working Area Size (W*D*H): Min. 1690*400*900mm - Exhaust Airflow Volume: 500-600m3/h - Material: 304 stainless steel - Standard Accessories: Spotlight; UV lamp; Fluorescent lamp; Exhaust system; 4m exhaust duct; Pipe strap; Auto flushing device; Stainless steel hanger; Adjustable hot and cold water tap; 180mm deep stainless steel water sink, funnel with filter. Active carbon filter - Functions: Odor-remove function. Sink, two-person operations. Exhaust air system: High-power suction device, low noise. Water heater. Pulverizer to avoid blocking drainage pipes. Automatic flushing device. |  |  |  |
| 2 | Refrigerator with glass door Quantity: 1 unit Minimal technical specification: Refrigerator with glass door, interior lighting, automatic defrost, mechanical temperature control, min. 5 height-adjustable shelves, wheels, ventilation cooling. External dimensions: min. 600 x 600 x 2025 mm Volume (gross): min. 385 lit. Temperature range: min. (+25°C,60% Rh) -2 ...+14 °C Consumpition: max. 2,14 kWh/24h Gas: R600A Volume (net): min. 360 lit. |  |  |  |
| 3 | Ventilated cabinet Quantity: 1 set Minimal technical specification: External dimensions (LxWxH): min. 1200x500x1990 mm Internal dimensions (LxWxH): min. 1100x450x1500 mm Total volume: min. 750l - Control panel with keyboard and LCD min. 4,3" color display where are displayed all process parameters. - Shelves made of stainless steel, adjustable in height, flanged and with basin. - Safety basin to collect any pouring from containers. Functions: - Automatic Suction Possibility - Function delay Possibility - Open door alarm |  |  |  |
| 4 | Laboratory tables Quantity: 4 units Minimal technical specification: - Dimensions (LxWxH): min. 140x65x80 cm - Worktop: compact laminate - Metal plastic coated stand - Frame: min. 40x40mm tube with adjustable feet |  |  |  |
| 5 | Writing table with drawers Quantity: 2 units Minimal technical specification: - Dimensions (LxWxH): min. 150x70x75 cm - Worktop formed melamine board min. 36mm - Metal plastic coated stand - Frame: min. 40x40mm tube with adjustable feet - All edges edged with ABS edging tape min. 2 mm thick |  |  |  |
| 6 | Drawers Quantity: 3 units Minimal technical specification: - Dimensions (LxWxH): min. 50x50x65 cm - Material double-sided melamine plat - 1 drawer, cupboard with shelf - Wheels fi 50 silicone, front with brake - All edges edged with ABS edging tape min. 2 mm thick |  |  |  |
| 7 | Laboratory chairs Quantity: 7 units Minimal technical specification: - Contoured back. - Gas height adjustment. - Height adjustable back. - Back angle adjustment. - Large 5 star base - Min.50mm twin wheel castors. - Seat height: min. 450 - 580mm - Seat depth: min. 400mm - Seat width: min. 450mm |  |  |  |

| All items | Supply, delivery and installation performed by contractor or authorised service provider. All the equipment must include all necessary parts and standards for its installation. |  |  |  |
| --- | --- | --- | --- | --- |
| All items | Testing of all basic functions of the instrument on a set of producers standard samples commonly used for the corresponding instrument. |  |  |  |
| All items | Technical documentation for equipment (Operating manuals/ Users Guide/ Equipment operating instructions/ Cleaning procedures/ Maintenance procedures/ Calibration procedures) upon delivery. Technical documentation with drawings for Aquatic weed harvester, Arms and Implements must be delivered for the purpose of aquatic weed harvester registration as technical floating object. |  |  |  |
| All items | - Warranty - Tenderers must provide local reliable warranty service agent providing maintenance and the rapid supply of equipment spare parts and consumables for the Warranty duration of one year. - Offer must include warranty service description including: - Service organisation contact data including name, postal address, telephone number, fax number and e-mail address; - Help Desk (phone) support, which must be available during working hours, 09.00 – 17.00 CET; - Guaranteed maximum response time to submitted maintenance support request (fax or e-mail) of 2 (two) working day; - Guaranteed that any requests for services will be attended to within 72 hours; - Guarantee that all items can be repaired or alternatively replaced within a maximum of 72 hours; - Guarantee that genuine spare parts and consumables will be available for a period of minimum 10 years from the date of final acceptance of the equipment. |  |  |  |
| All items | Commercial warranty Min. 1 year from the date of delivery or min. 1000 hours of operation in accordance with the conditions laid down in Article 32 of the General Conditions and Article 33 of the Special Conditions. Detailed description of the organisation of the proposed service and description of the Manufacturer’s commercial warranty shall be included in the offer. |  |  |  |