# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Procurement of medical equipment for project "Romania-Serbia joint initiative against cancer in cross-border region: improved diagnosis and treatment of malignant tumors - ROSECAN"

Publication reference: RORS193-LB-TD7

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Intensive and Half Intensive Care Monitor with Central Station Within Each Intensive Care Unit (ICU) Quantity: 14 unit 1.1 Patient monitors with monitoring of IBP and EtCO2 (4 pcs) Technical requirement: - Patient monitor for patients of all ages (adult, children, neonates). - Monitor has color display, minimum 12“ diameter, resolution min 1280x800, touchscreen - Viewing angle at least 170º. - Displaying minimum 10 waveforms - Automatic screen brightness adaptation to ambient light. - Possibility to set at least 15 quick softkeys on screen that are frequently used, for easy access to needed functions. Possibility to set at least 3 timers on screen - Trend memory: minimum 240 hours with resolution of 1 minute, min 2000 events, min 48 hours full disclosure , 3000 NIBP results - Integrated ports: 2 x USB, network (RJ45). - Monitor has active functions for following measurements: - 3/5 ECG - ST analysis. - QT analysis - Minimum 24 arrhythmia analysis. - Respiration - NIBP – non-invasive pressure - SpO2 – saturation - Temperature - Monitoring IBP – 2 channels - Monitoring EtCO2 - Monitor has Early Warning Score function with dynamic refresh of scoring according to change of vital signs. - Users can view mini trend panel on screen - Manuel event function for user to quickly capture screen snapshot including all current numeric and waveforms. - Modules and accessories that must be delivered with monitors - ECG cable 5-lead – 4 pcs - NIBP tube and cuff – 4 pcs - SpO2 cable and sensor – 4 pcs - Temperature probe – 4 pcs - Cables and transducers for 2-channel IBP – 4 sets - Plug & Play module for EtCO2 - 2 pcs - Accessories for monitoring EtCO2 – 2 sets - Rail mount system Plug & Play module for EtCO2 - 2 pcs - Accessories for monitoring EtCO2 – 2 sets - Rail mount system 1.2 Patient monitors with monitoring of IBP (12 pcs) Technical requirement: - Patient monitor for patients of all ages (adult, children, neonates). - Monitor has LED color display, touchscreen, minimum 12“ diameter, resolution at least 800x600 - Displaying minimum 11 waveforms - Support review and storage of 120 hours trend (resolution 1 min), 1800 alarm events, 1600 NIBP measurements, 128 ARR alarms and 48-hour full disclosure waveforms - Integrated ports: 2 x USB, network (RJ45). - Data transfer via USB drive - AC and battery power. At least 4 hours of working time with large capacity Li-ion battery - Design without fan inside of housing - Monitor has active functions for following measurements: - 3/5 ECG - Multi-lead ECG algorithm - ST analysis. - QT analysis - Minimum 24 arrhythmia detection - Respiration - NIBP – non-invasive pressure - SpO2 – saturation - Temperature - Monitoring IBP – 2 channels - Accessories that must be delivered with monitors - ECG cable 5-lead – 12 pcs - NIBP tube and cuff – 12 pcs - SpO2 cable and sensor – 12 pcs - Temperature probe – 12 pcs - Cables and transducers for 2-channel IBP – 12 sets - Rail mount system 1.3 Central Monitoring Station (1 pcs) Technical requirement: - Complete network for patient monitoring. System has 1 PC , 2 displays at least 24 inch, 1 keyboard and 1 mouse. - Minimum 16 patient monitors can be connected to one station - Displayed waveforms: ECG, Pleth, Resp, CO2, IBP - Displayed numerics: HR, ST, PVCS, QT/QTc, RR, SpO2, PR, NIBP, TEMP, IBP, CO, EtCO2 - Audible and visual alarms with high, medium and low priority - minimum 240 hours of tabular and graphic trends for all parameters - minimum 240 hours of full-disclosure waveforms - 1000 events with parameter name and 16-second waveform before and after triggering alarm - At least 8 hours of minitrends for all parameters - Tools: hemodynamic calculations, oxygenation calculation, ST histogram, hemographic data |  |  |  |
| 2 | Operating tables Quantity: 4 units Technical requirement: - Operating table for General surgery, Urology and Gynaecology - Mobile operating table with minimum 5 segments - Electric motorized adjustment of table movements - Integrated battery supply with table autonomy, 2 rechargeable batteries. Battery duration up to 5 working days . - Minimal maximum table load 400 kg in central position of column and table top Adjustment of operating table - Trendelenburg positioning min :+25⁰/-25⁰. - Lateral tilt min 25⁰ on both sides. - Longitudinal shift min. 330 mm. - Motorised back section in minimal range +80⁰/-60⁰ - Table width with rails minimum 600mm Standard accessories for mobile table - Leg section, two part spredable with pad - Head section, single joint with pad - Anaesthesia screen extendable, height adjustable via clamp - Clamp radial setting - Clamp basic - Infusion pole, 4 hooks, with clamp - Body strap, with locking clamps - Arm support with clamp - Universal and hand cuff - Pelvis extention with pad - Goepel leg holder, pair - Upper back section, with pad - Cable remote control - Drainage bowl, made of stainless teel - Instruemet tray, suitable for drainage bowl adapters - Drainage bowl adapter, made of stainless teel |  |  |  |
| 3 | Operating Reflector for the OR Quantity: 1 unit Technical requirement: - Main and satellite light ceiling, LED - High reliability with at least 80 separate LED lights - Main light + Satellite - Illumination of the main light minimum l60.000 lx - Illumination of the satellite light minimum l40.000 lx - Color rendering for main light : - index Ra: min.93 - index (blood) R9: min.80 - index (skin): R13 min. 94 - Pattern size variable between 0,8m-1,2m : 17-25cm - Average service life of the LEDs more than 60.000 h - Color temperature fixed, in range 4000-5000 K |  |  |  |
| 4 | Operating Reflector for OR with FHD camera, receiver and monitor Quantity: 1 unit Technical requirement: - Main and satellite light ceiling, LED - Illumination of the main light minimum l60.000 lx - Illumination of the satellite light minimum l40.000 lx - Touch screen remote control for surgical lights. - Color temperature control 3500 K – 5000 K in minimum 3 steps. - Tablet control of all OR light functions - Light intensity 160.000 lx from 0,8 m to 1,3 m distance from operation field. - Light intensity of 160.000 lx on pattern size 16 - 25 cm. - Colour rendering ( index ): Ra min.97 - Wireless HD camera in centre of light (main or satellite). Main and satellite lights need to be HD camera ready. - Wireless receiver - Full HD medical monitor min. 24" on separate arm. |  |  |  |
| 5 | Automatic Sterilizer - Autoclave Quantity: 1 unit Technical requirement: Chamber volume min. 60 liters, 1 STU, single door (non-pass-through model) Purpose / use: The device is used for sterilization of various medical materials with steam at temperatures of 121-134 ° C. Medical material (packaged or unpackaged) that is to be sterilized is placed in baskets and boxes, and these are then placed in an internal carrier in the sterilizer chamber. Description: Microprocessor controlled: steam sterilizer, with manual vertically sliding door (one) - non-pass-through model. The device is designed as a cabinet model with stainless steel side panels. Steam supply from own built-in steam generator. Chamber shape: chamber with square cross-section and rounded edges (round chamber not acceptable) Chamber volume: min. 60 lit. (1 STU) The chamber must be thermally insulated. The chamber door is manually operated, sliding vertically, heat and sound insulated. The door must have a security system to prevent it from opening during the process. Electrical connection: 3 phases, 400V, 50 Hz Installed device power: 6 kW (steam generator power: max. 4 kW) The steam generator heaters are not in direct contact with water Feed water: municipality water Design pressure of pressure vessels: Full vacuum and min. 3.5 bar overpressure. The sterilizer is equipped with a software combination: - Standard steam sterilization processes at 120-135 ºC (textiles, instruments, rubber, plastics) - Express cycle for sterilization of a very small amount of unpackaged instruments at 134 ºC. - Test programs: "Bowie & Dick", leak test, etc. The device has an injector vacuum pump. A modern processor unit enables the adjustment of certain parameters, computer control of a PC or Laptop, modification of existing sterilization programs, creation of new ones, etc. The device is controlled via the control panel which is located on the front of the device, above the chamber. The control panel is "touch-screen" type, with a minimum diagonal of 8 ". The device has audio and visual signals to indicate the beginning / process / end of the process and process error. The device itself has a thermo-printer (alpha-numeric) that records the parameters of the sterilization process. The sterilizer has a series of safety measures to prevent the door from opening during sterilization, a safety valve, measuring and control elements, sound and light alarms, etc. Material: Chamber material: AISI 316L Device formwork material: AISI 304 Pipe installation and valve material: copper, brass and stainless steel. Dimensions: Appliance dimensions: max. 600x1600x900 mm (width x height x depth) Nominal chamber dimensions: min. 300x300x600 (width x height x depth) Standards: - EN ISO 13485 - quality system in the production of medical devices - MDD (Medical Devices Directive) 93/42 EEC - European Directive for Medical Devices - PED (Pressure Equipment Directive) 97/23 EC - European Directive for pressure vessels - (SRPS) EN 285 - European standard for large sterilizers Accessories: - A water purification system is supplied with the appliance as an accessory to prepare for the steam generator purified water of appropriate capacity and quality. |  |  |  |
| 6 | Mechanical Hospital Beds with mattresses Quantity: 100 units Technical requirement: - Total bed dimensions: 900mm x 2050mm ±20mm - Working load (SWL) min. 150 kg - Bed frame is made of steel profiles painted with powder-colored polyester - Fixed lying plate, made from steel wire - Mechanical bed, with possibility to adjust head section with crank - Possibility to install patient lifting pole - Four Castors, min Ø125mm |  |  |  |
| 7 | Nightstands Quantity: 100 units Technical requirement: - Patient locker with two shelfs and one drawer, on castors with breaks - Locker dimensions 480mm x 450mm x 750 ± 30mm - Max. Load on locker top 20kg, shelfs and drawer 5kg |  |  |  |
| 8 | Electric Beds for Intensive care Unit Quantity: 14 units 8.1 ICU Hospital Bed with four fences (4 pcs) Technical requirement: - Overall dimensions 990mm x 2200mm ±10mm - Integrated bed extension minimum 280 mm - Castors Ø150mm, with central locking system. - Regulation of height with electric motor in min. range: 380mm - 840mm +- 10mm - Electrical automatic Tr / ATr ≥ -14 °./ +14° - Four-part safety sides, one hand operation, made of quality polymer with integrated control panels on two safety sides of the bed. - Lead gel battery for starting all electrical functions when the bed is disconnected from the main supply - Infusion stand with four hooks - Patient lifting pole with adjustable grip 4 deflection rollers at the corners of the bed - Oxygen tank holder - Anti-Decubitus mattress with elastic PU cover, following dimensions and features of the mattress base 870mm x 2000mm±10mm - Maximum weight of patient min. 215 kg 8.2 ICU Hospital Bed with two fences (10 pcs) Technical requirement: - Overall dimensions 990mm x 2150mm ±10mm Bed extension minimum 280 mm - Castors Ø150mm, with central locking system. - Regulation of height with electric motor in min. range: 380mm - 800mm +- 10mm - Electrical automatic Tr / ATr ≥ -14 °./ +14° - Patient handset control unit - One part safety sides covering large proportion of the mattress base, with easy swivel down and rising up. - Infusion stand with four hooks - Patient lifting pole with adjustable grip 4 deflection rollers at the corners of the bed - Anti-Decubitus mattress with elastic PU cover, following dimensions and features of the mattress base 870mm x 2000mm±10mm - Maximum weight of patient min. 215 kg |  |  |  |

| All items | Installation performed by contractor or authorised service provider. All the equipment must include all necessary parts and standards for its installation. |  |  |  |
| --- | --- | --- | --- | --- |
| All items | Testing of all basic functions of the instrument on a set of producers standard samples commonly used for the corresponding instrument. Installed equipment must be tested as system, compatible with existing system. |  |  |  |
| All items | Technical documentation for equipment (Operating manuals/ Users Guide/ Equipment operating instructions/ Cleaning procedures/ Maintenance procedures/ Calibration procedures) upon delivery. |  |  |  |
| All items | Warranty Tenderers must provide local reliable warranty service agent providing maintenance and the rapid supply of equipment spare parts and consumables for the Warranty duration of one year. Offer must include warranty service description including: - Service organisation contact data including name, postal address, telephone number, fax number and e-mail address; - Help Desk (phone) support, which must be available during working hours, 8AM – 6PM; - Guaranteed maximum response time to submitted maintenance support request (fax or e-mail) of 1 (one) working day; - Guaranteed that any requests for services will be attended to within 24 hours; - Guarantee that all items can be repaired or alternatively replaced within a maximum of 72 hours; - Guarantee that genuine spare parts and consumables will be available for a period of minimum 3 years from the date of final acceptance of the equipment. |  |  |  |
| All items | Commercial warranty 1 year (after the end of 1 year standard warranty) in accordance with the conditions laid down in Article 32 of the General Conditions and Article 33 of the Special Conditions. Detailed description of the organisation of the proposed service and description of the Manufacturer’s commercial warranty shall be included in the offer. |  |  |  |