# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Supply of Procurement of Liquid chromatography–mass spectrometry instrument	p 1 /…

Publication reference : 2021/S 213-563449

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Liquid chromatography–mass spectrometry instrument 1 set consisting of: Pump - 1pcs - Quaternary pump - Dual pistons in series pump with proprietary servo-controlled variable stroke design - Gradient formation: low pressure quaternary mixing - Active seal wash - Automatic purge valve - Automatic compressibility compensation - Flow range: 0.001 - 5 ml/min, (or wider) - Pressure: up to 700 bar or better - Integrated degassing unit with 4 channels - Flow precision: ≤0.07% RSD or better - Flow accuracy: ±1% or better - Composition precision: ˂0.15% RSD or better - Composition accuracy: ±0.5% or better Autosampler -1pcs - Injection range: 0.1 – 20 µl in 0.1 ul increments or wider/better - Injection precision: ˂ 0.25% RSD or better - Pressure range: up to 0-700 bar or better - Sample capacity: 100 vials x 1,5 ml or better - Carry over: 0.004% or better - Injection cycle time: 10 sec or better Column termostat – 1pcs - Temperture range :+5°C to 85°C or better - Temperature stability: ±0.05 °C or better - Temperature precision: 0.1 °C or better - Heat up time: 15 min from 20 to 40 °C or better - Cool down time: 15 min from 40 °C to 20 °C or better - Column capacity: min 2 column of 300 mm length UV detector diode array - 1pcs - Light source deuterium and/or tungsten lamps - Detector type: 1024-element diode array - Wavelength range: 190-640nm or wider - Wavelength accuracy: ±1 nm, self-calibration with deuterium lines - Wavelength precision: ±0.1nm or better - Number of signals : min 8 - Short-term noise: ±6·µAU or better - Drift: <1mAU/h - Data rate:125 Hz - Flow cell 10mm Mass detector - 1pcs - Triple quadropole - Ionization sources : Electrospray(ESI) source, and Atmospheric pressure chemical ionization (APCI) source - Mass range : 5 to 3000 m/z or better - Mass resolution: 0.7 Da - Mass stability: ˂ 0.1Da/24h - Dynamic range: ˃10⁶ - Scan rate : 15000 Da/sec or better - Polarity switching : 25 msec or better - MRM sensitivity S/N ESI positive : 1 pg reserpine injected on column; quantifying on the transitionm/z 609.3 – 195.1 S/N ≥200000:1 or better - MRM sensitivity S/N ESI negative: 1 pg chloramphenicol injected on column; quantifying on the transition m/z 321.0 – 152.0 S/N ≥200000:1 or better - Dwell time: 0.5 msec - Maximum MRM aquisition rate: 500 MRM s/sec or better - MRM transitions: 450 per time segment, 13500 ion transitions per method or better - Detector type: High energy conversion dynode with multiplier or discrete-dynode detector - Vacuum system: One turbomolecular pump with one mechanical pump or better - Triple quadropole system:Must be able to perform simple maintenance without need to break vacuum. - Triple quadropole system: Must have automatic optimization of ion optics, and calibration of masses with internal tune solution, software controlled Computer with operating system and printer – 1 set - Brand name PC with operating system, monitor and printer according to the manufacturer's recommendation Software -1pcs - Software for data collection, processing and reporting of results, control of entire system parameters, data processing in qualitative and quantitative terms with all databases and suplementary modules (if necessary) necesary to process samples (at least water and soil) containing pesticides, including neonicotinoids, insecticides, herbicides and other chemicals commonly used in agriculture. Nitrogen generator - 1pcs - Nitrogen generator with built-in air compressor, capacity 30 L / min or better, or in according to the recommendations of the mass detector manufacturer Set of consumables necessary for the installation of the instrument - 1 set - Fittings, capillaries, UPLC columns, pre-columns if needed, and other instrument consumables, if necessary, for analyzing and quantifying pesticides in 2000 samples. Set of tools necessary for the maintenance of the instrument - 1 set - Recommended by the equipment manufacturer Instalation and training - Instalation and commision of system in premises of Contracting Authority is obligation of supplier - Training to be performed for up to 5 persons designated by Contracting Authority is obligation of supplier |  |  |  |