# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Supply of medical instruments, tools and equipment	p 1 /12

Publication reference : RORS 284 / Municipality of Jimbolia / 03 – RE2

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation (Please mention supporting document and page from this document where specification/s is/are included) | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Lot 8 - EKG machine -2 pieces |  |  |  |
|  | - GENERAL CHARACTERISTICS |  |  |  |
|  | - QUALITY STANDARDS, CONFORMITY AND SAFETY |  |  |  |
|  | The equipment must meet all legal conditions for placing on the market in Romania: |  |  |  |
|  | - Medical Devices Directive 93/42/EEC (medical equipment directive, CE marking for Europe) The CE Certificate as well as the CE Declaration of Conformity will be presented |  |  |  |
|  | - ISO 9001 Certificate and ISO 13485 certificate for the producer |  |  |  |
|  | - EN ISO 9001:2008 quality management system for the seller |  |  |  |
|  | - The SERVICE representative of the manufacturer in Romania must have implemented an ISO 9001 certified quality management system regarding the repair of this equipment (to attach proof of this). |  |  |  |
|  | - The technical staff (service specialists) must be trained by the manufacturer and certified by him (a proof to this effect must be attached). |  |  |  |
|  | - EQUIPMENT GENERAL DESCRIPTION |  |  |  |
|  | - Wireless medical tablet system |  |  |  |
|  | Medical tablet with wireless connection possibility |  |  |  |
|  | Wireless and portable ECG tablet, a powerful diagnostic device, to make it easy to capture, analyse and share high-quality ECG data that can be shared and accessed for consultation with cardiologists or specialists from any device. |  |  |  |
|  | ECG data should be available in real-time, on the tablet and on the physician's screen, without waiting for the ECG printout. An immediate access to connected tablet of ECG should reduce the administrative workload and have powerful ECG analysis and data connectivity always at hand. The tablet should have installed application consultation to make it simple, safer and every time. The physicians and specialists must be able to request a repeat ECG even before the electrodes are removed from the patient. This gives valuable time back to the patient care. |  |  |  |
|  | The high resolution ECG signal must have posibility to be saved in a digital format, enabling the possibility of reviewing and managing the ECG data later on. |  |  |  |
|  | Physician or a specialist who is asked for consultation can access the ECG recording from any device or platform and have exactly the same, advanced interpretation possibilities as the physician asking for consultation. This way, wrong, unnecessary referrals to specialist and future, costly treatments can be reduced. |  |  |  |
|  | Operating sistem : open source software |  |  |  |
|  | Procesor : CPU at least 1.0 GHz per core |  |  |  |
|  | Barcode reader: at least 1D/2D barcode imager |  |  |  |
|  | Screen: min 1280 x 720 px IPS |  |  |  |
|  | Camera: minimum 5 MP |  |  |  |
|  | Memory: at least 5 GB Connectivity: minimum Wi-Fi 802.11b/g/n and 2.4 GHz single band Bluetooth 4.1 |  |  |  |
|  | Security: 2 step authentication - User password or PIN, SMS verification code HTTPS transmission protocol |  |  |  |
|  | Docking station specifications |  |  |  |
|  | Conectivity: minimum Wi-Fi 802.11b/g/n, 150 Mb/s |  |  |  |
|  | Frequency: 2402-2480 MHz; Band width: 40 Mhz |  |  |  |
|  | Battery and supply: AC/DC adaptor Input: 100-240 V ; 50-60 HZ; 600-300 mA Output: 5 Vdc, 5000 mA Battery type: Lithium-Polymer rechargable Minimum capacity 8800 mAh Operating time: up to 8 hours |  |  |  |
|  | Operating conditions: - temperature : 10-40 Celsius degrees humidity : 25-85% |  |  |  |
|  | 2. ECG system |  |  |  |
|  | Advanced patient data system, with set of powerful analysis tools and application to enable viewing of high-quality ECG data and advanced ECG analysis from any web-enabled device, from any location and anytime. |  |  |  |
|  | Physician can share and consult about the ECG recording in a safe and simple way, with just a few clicks, while patient data stays secure. |  |  |  |
|  | Physician or a specialist who is asked for consultation can access the ECG recording from any device or platform and has exactly the same, advanced interpretation possibilities as the physician asking for consultation. This way, wrong, unnecessary referrals to specialist and future, costly treatments can be reduced. |  |  |  |
|  | All data can be wireless syncronised with the existing data of the same patient in the hospital system through HL7 standard |  |  |  |
|  | Patient data: name of the doctor and the patient, heart rate, print speed, sensitivity, electrodes check, filters, interpretation, measurements history, result opinions, date and hour of the measurement |  |  |  |
|  | Possibility of wireless printing of the results direct to a pirnter, wireless |  |  |  |
|  | Possibility to attach to the patient file a picture, using nothing elese than the tablet |  |  |  |
|  | Leads display: minimum 6x6+1, 6x6, 3x3 auxiliar, 6 main, 6 auxiliar, 3x4, 12x1 |  |  |  |
|  | Safety standards: IEC/EN 60601-1; IEC/EN 60601-2-25; IEC/EN 60601-1-2 (EMC) |  |  |  |
|  | Size: minimum 10.5 cm × 4 cm × 4.2 cm |  |  |  |
|  | Lowpass filters: minim 150 Hz, 250 Hz Highpass filters: minim 0.05 Hz, 0.2 Hz, 0.5 Hz Mains filter: 50/60 Hz Myogram (muscle tremor) filter: 25 Hz (40 dB/dec) or 35 Hz (20 dB/dec) |  |  |  |
|  | Sampling frequency: mimim 32000 sample / secund / channel Frequency of ECG analysis: minim 1000 sample / secund |  |  |  |
|  | Detection of loose electrodes Sampling frequency: at least 32.000 samples/second/channel |  |  |  |
|  | ECG analysis frequency: minimum 1.000 samples/second CMRR: > 110dB |  |  |  |
|  | Input impedance: 10 Mohm Digital resolution: 2.5μV / 19-bit |  |  |  |
|  | Pacemaker detection: ≥ ± 2 mV/ ≥ ± 0.1 ms Heart rate measurement: up to 300 BPM |  |  |  |
|  | - Spirometry module |  |  |  |
|  | Multiple measurement modes and parameter calculation options: quick, primary or advanced spirometry app package |  |  |  |
|  | Automatic selection of the best of all repeated maneuvers |  |  |  |
|  | Detailed report with the ability to switch between charts and values for a clear interpretation |  |  |  |
|  | Option to compare different measurements on the same screen |  |  |  |
|  | Calibration: Automated self-calibration based on the ambiental conditions and pre-calibrated mouthpieces |  |  |  |
|  | Accuracy & Reproducibility: Meets or exceeds ATS (1994), ERS (1993), and ATS/ERS (2005) |  |  |  |
|  | Measurement range: Volume range: 0-14 L; Flow range: +/- 14L/ sec |  |  |  |
|  | Quality checks: ATS Acceptability and ATS Reproducibility checks |  |  |  |
|  | - PACKAGING, SHIPPING AND TRANSPORTATION |  |  |  |
|  | The transport of medical equipment to the unit's headquarters is the responsibility of the supplier |  |  |  |
|  | The supplier has the obligation to properly pack the device so that it can withstand handling during transport and arrive in good condition at the final destination |  |  |  |
|  | - TERMS OF DELIVERY |  |  |  |
|  | The delivery will be made at the headquarters of the beneficiary unit. |  |  |  |
|  | The CE Declaration of Conformity of the manufacturer will be presented. |  |  |  |
|  | The device must have an original user manual |  |  |  |
|  | - RECEIPT OF THE PRODUCT |  |  |  |
|  | At the reception of the product, the presence of at least one authorized representative of the supplier is mandatory |  |  |  |
|  | Reception will take place at the final destination. If the inspected product does not meet the specifications or is damaged, the contracting authority has the right to reject it, and the supplier has the obligation to replace the non-compliant products. |  |  |  |
|  | - PERSONAL TRAINING |  |  |  |
|  | The training of the medical personnel is done free of charge, at the place of operation of the device |  |  |  |
|  | - SERVICE, WARRANTY AND POST WARRANTY |  |  |  |
|  | Full guarantee for at least 24 months from commissioning, a statement to this effect will be attached |  |  |  |
|  | During the warranty period, all costs related to the repair and/or replacement of damaged devices, accessories or components are carried out free of charge by the device supplier, according to the conditions of the warranty certificate |  |  |  |
|  | Supplying with spare parts and consumables throughout the warranty and post-warranty period, at least 8 years from the date of delivery |  |  |  |
|  | The delivery to the assembly site with all the operations that this entails (transportation, devices for lifting and handling the equipment), as well as the installation, assembly, commissioning of the equipment, training of the operating personnel, will be ensured by the supplier with authorized personnel and operations are included in the price. A statement to this effect will be attached. |  |  |  |
|  | The response time to service requests during the warranty and post-warranty period will be a maximum of 48 hours from the receipt of notification from the beneficiary, |  |  |  |
|  | - COMMISSIONING |  |  |  |
|  | After commissioning and performance of functional tests, the "Installation and Commissioning Minutes" will be drawn up, signed by both parties. The training of the medical staff will be done by the provider's specialists |  |  |  |

Tenderers shall draw up the technical specifications in an organized manner, so that it may provide an easy check of compliance with the requirements/specifications provided under this section.

- Technical specifications indicating a particular origin, source, production, a factory or trade mark, an invention patent, a manufacturing license are listed in the specifications for the sole purpose of an easy identification of the product type and are not aimed at favouring or eliminating any economic operators or certain specific products, and such specifications shall be deemed to be provided with the note “OR EQUIVALENT”

- Any statement in the technical specifications regarding national standards / systems, national certificates / certifications, national references, national indicators, will be interpreted "or equivalent" for international entities;

- The requirements in this document shall be deemed minimal and compulsory to be met (all technical features must be met). To this end, the basic tender submitted, with deviations from the provisions of the tender offer, shall be considered only to the extent that the technical proposal entails the provision of a quality level higher than that of the listed requirements.

- Any technical features exceeding or above those required by the contracting authority shall be listed in the notes column and indication shall be provided of the page number in the technical documentation where the feature is listed.

- If, throughout the progress of the contract, it is found that elements of the technical proposal are non-compliant (inferior to or inconsistent with the requirements provided in the tender offer), the contracting authority/entity reserves the right to terminate the contract for convenience, or to claim a suspension in the supply of products until the situation detected has been remedied.

- Tendered products shall be new and not used.

- all equipment must be CE marked;

CONTRACTING CONDITIONS:

Delivery documents:

a) Invoice,

b) Shipping note, if the case

c) Quality certificate,

d) Technical book (manual) - in Romanian and in English

e) Certificate of warranty,

f) Certificate of conformity.

Inspection of the supplier products shall be carried out both quantitatively and qualitatively.

The qualitative inspection shall consist of inspecting the quality certificate or the declaration of conformity. The quantitative inspection shall consist of the inventory of delivery products.

Upon delivery, the parties shall execute the following:

a) The Delivery-Receipt Protocol (quantitative and qualitative acceptance);

b) The Commissioning Protocol;

c) The Staff Training Protocol.

Packaging and carriage:

Packaging shall be carried out according to product packaging instructions, in such a way so that they are protected during carriage and storage.

Products shall be delivered and assembled, so as to fall within the deadlines provided in the agreement.

Supplier obligations and responsibilities

The supplier undertakes to:

- Supply the product/s under the terms set in the contract

- Provide the supply, delivery and commissioning, according to all legal requirements applicable in Romania;

- Provide the enforcement of technological and organisational steps leading to the strict compliance with the technical specifications;

- Observe compulsory provisions relating to occupational health and safety, fire prevention and extinguishing, and environmental protection, during the delivery of products

- Not impact the beneficiary’s property, the environment or other properties pursuant to the delivery/handling;

- Notify the beneficiary, via phone or fax, on the date when the delivery and the inspection are to take place, at the location indicated by the contracting Authority;

- Replace, during the grace period, all inadequate components, from a qualitative point of view, such components rejected by the beneficiary, within a maximum of 15 days from communication thereof;

- Remedy any failure occurring during the warranty period, at their own expenses, under the terms set in the agreement, within a maximum of 48 hours from the time of notification thereof.

- Provide the delivery and setup of the equipment, at the headquarters of the Dr. Karl Diel Hospital of Jimbolia, in the town of Jimbolia, Str. Contele Csekonics, nr. 4

- Install all delivered equipment, according to the beneficiary’s requirements

Obligations of the contracting authority

The contracting authority undertakes to:

- Indicate the location where the product is to be delivered and accepted under a Protocol;

- Provide access to the location and utilities necessary for commissioning and operating the products;

- Provide staff for carrying out the quantitative and qualitative acceptance;

- Notify the supplier in writing on any failure occurred during the warranty period

- Pay the price for the product, according to the executed agreement and within the payment deadline set.