# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Supply of medical instruments, tools and equipment	p 1 /8

Publication reference : RORS 284 / Municipality of Jimbolia / 03 – RE4

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation (Please mention supporting document and page from this document where specification/s is/are included) | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Lot 1 - ORTHOPEDIC DRILL-REAMER - 1 piece |  |  |  |
|  | DESCRIPTION Cordless, battery powered rotary drill for use during general surgical procedures |  |  |  |
|  | TECHNICAL CHARACTERISTICS OF THE DRILL |  |  |  |
|  | Pistol shape with two triggers and a function switch that can lock one or both triggers. |  |  |  |
|  | two directions of rotation operated by buttons |  |  |  |
|  | Each trigger separately controls the variable speed operation of the hand piece, either clockwise or counterclockwise. |  |  |  |
|  | When operated at the same time, the triggers trigger the “oscillating” operating mode for inserting separate pins: forward and backward |  |  |  |
|  | The switch blocks one or both triggers in order to prevent the accidental and improper operation of the drill |  |  |  |
|  | It is provided with the “shift collar” - speed and torque shift lever, with the help of which the speed and torque can be set; - Drilling - counter clockwise rotation: high speed and low torque Boring - clockwise rotation: low speed and high torque |  |  |  |
|  | cannulated on the entire working length |  |  |  |
|  | The battery can be detached by actuating a safety button |  |  |  |
|  | Variable rotation speed: The maximum rotation speed when drilling must be min. 1200 rpm, and for boring min. 270 rpm. |  |  |  |
|  | It must have min. 3 adapters for drilling and boring : Dual trigger pin collet 0.7-2 mm -1unit Dual triger large collet 2.0-3.2 mm-1 unit Large reamer attachement : type AO- 1 unit These adapters must be autoclavable (132 – 134° C) |  |  |  |
|  | The adapters must be attached and detached by activating the “shift collar” |  |  |  |
|  | Maximum weight: 1.10 kg |  |  |  |
|  | It must withstand disinfection and sterilization processes: class Ipx9 |  |  |  |
|  | BATTERIES |  |  |  |
|  | Autoclavable (132 – 134° C) |  |  |  |
|  | Rechargeable |  |  |  |
|  | Lithium-Ion |  |  |  |
|  | Output voltage: 9.9 V DC |  |  |  |
|  | Easy coupling/uncoupling to/from the drill |  |  |  |
|  | Safety button for connecting/disengaging the oscillating and boring bits |  |  |  |
|  | In must send to the charger information on the number of uses and the current power level |  |  |  |
|  | It must be equipped with a light indicator for displaying the charging/discharging status of the battery; |  |  |  |
|  | It must have a light indicator showing when the battery is discharged. |  |  |  |
|  | When the indicator light is flashing, the load on the hand piece is decreased so as the battery life is extended. |  |  |  |
|  | Integrated microprocessor for storing the operating data of the connected drill for remote diagnosis when it is docked in the charger and it is connected to the Internet via the network; |  |  |  |
|  | Maximum weight: 0.30 kg - small battery; 0.45 kg - large battery |  |  |  |
|  | CONNECTION ADAPTER (CHUCK) FOR VARIOUS TYPES OF COUPLING |  |  |  |
|  | Autoclavable (132 – 134° C) |  |  |  |
|  | Available for the coupling type: Jackobs with key |  |  |  |
|  | Coupling to the drill by means of a button, without an auxiliary key |  |  |  |
|  | DELIVERY TERM |  |  |  |
|  | The equipment shall be delivered no later than 20.12.2023. The equipment shall be delivered at the headquarters of Dr Karl Diel Hospital in Jimbolia (Contele Csekonics Street no.4). The shipping costs are included in the price of the tender. The supplier shall fill in an affidavit related to meeting these requirements |  |  |  |
|  | WARRANTY PERIOD |  |  |  |
|  | The supplier shall grant a minimum 24-month warranty period from the date of commissioning The response time in case of malfunctions and/or breakdowns of the equipment = max 48 hours from the notification The supplier shall carry out at least one technical check of the equipment and issue a check report during the warranty period The supplier shall fill in an affidavit related to meeting these requirements |  |  |  |
|  | TRAINING AND COMMISSIONING |  |  |  |
|  | The supplier shall ensure the training of the users of equipment with qualified staff (an affidavit must be submitted related to meeting these requirements) |  |  |  |
|  | The supplier shall ensure the commissioning of the equipment (an affidavit must be submitted related to meeting these requirements) |  |  |  |
|  | The supplier holds an operating license issued by the Ministry of Health for licensing of the supply, commissioning, repair and maintenance activities related to the tendered equipment |  |  |  |

Tenderers shall draw up the technical specifications in an organized manner, so that it may provide an easy check of compliance with the requirements/specifications provided under this section.

- Technical specifications indicating a particular origin, source, production, a factory or trade mark, an invention patent, a manufacturing license are listed in the specifications for the sole purpose of an easy identification of the product type and are not aimed at favouring or eliminating any economic operators or certain specific products, and such specifications shall be deemed to be provided with the note “OR EQUIVALENT”

- Any statement in the technical specifications regarding national standards / systems, national certificates / certifications, national references, national indicators, will be interpreted "or equivalent" for international entities;

- The requirements in this document shall be deemed minimal and compulsory to be met (all technical features must be met). To this end, the basic tender submitted, with deviations from the provisions of the tender offer, shall be considered only to the extent that the technical proposal entails the provision of a quality level higher than that of the listed requirements.

- Any technical features exceeding or above those required by the contracting authority shall be listed in the notes column and indication shall be provided of the page number in the technical documentation where the feature is listed.

- If, throughout the progress of the contract, it is found that elements of the technical proposal are non-compliant (inferior to or inconsistent with the requirements provided in the tender offer), the contracting authority/entity reserves the right to terminate the contract for convenience, or to claim a suspension in the supply of products until the situation detected has been remedied.

- Tendered products shall be new and not used.

- all equipment must be CE marked;

CONTRACTING CONDITIONS:

Delivery documents:

a) Invoice,

b) Shipping note, if the case

c) Quality certificate,

d) Technical book (manual) - in Romanian and in English

e) Certificate of warranty,

f) Certificate of conformity.

Inspection of the supplier products shall be carried out both quantitatively and qualitatively.

The qualitative inspection shall consist of inspecting the quality certificate or the declaration of conformity. The quantitative inspection shall consist of the inventory of delivery products.

Upon delivery, the parties shall execute the following:

a) The Delivery-Receipt Protocol (quantitative and qualitative acceptance);

b) The Commissioning Protocol;

c) The Staff Training Protocol.

Packaging and carriage:

Packaging shall be carried out according to product packaging instructions, in such a way so that they are protected during carriage and storage.

Products shall be delivered and assembled, so as to fall within the deadlines provided in the agreement.

Supplier obligations and responsibilities

The supplier undertakes to:

- Supply the product/s under the terms set in the contract

- Provide the supply, delivery and commissioning, according to all legal requirements applicable in Romania;

- Provide the enforcement of technological and organisational steps leading to the strict compliance with the technical specifications;

- Observe compulsory provisions relating to occupational health and safety, fire prevention and extinguishing, and environmental protection, during the delivery of products

- Not impact the beneficiary’s property, the environment or other properties pursuant to the delivery/handling;

- Notify the beneficiary, via phone or fax, on the date when the delivery and the inspection are to take place, at the location indicated by the contracting Authority;

- Replace, during the grace period, all inadequate components, from a qualitative point of view, such components rejected by the beneficiary, within a maximum of 15 days from communication thereof;

- Remedy any failure occurring during the warranty period, at their own expenses, under the terms set in the agreement, within a maximum of 48 hours from the time of notification thereof.

- Provide the delivery and setup of the equipment, at the headquarters of the Dr. Karl Diel Hospital of Jimbolia, in the town of Jimbolia, Str. Contele Csekonics, nr. 4

- Install all delivered equipment, according to the beneficiary’s requirements

Obligations of the contracting authority

The contracting authority undertakes to:

- Indicate the location where the product is to be delivered and accepted under a Protocol;

- Provide access to the location and utilities necessary for commissioning and operating the products;

- Provide staff for carrying out the quantitative and qualitative acceptance;

- Notify the supplier in writing on any failure occurred during the warranty period

- Pay the price for the product, according to the executed agreement and within the payment deadline set.